## Slide 1: Hardware and Interface PCB block diagram

## Slide 2
DC Accuracy +/-5mV @10V
Motor Controller Not Shown
HW and Interface PCB Prototype Schematic and Assy Diagrams

## Slide 3
Reed Switch, open->close step response

## Slide 4
100%
30%
Waiting
7%
Tank Level step response to 3-speed servo motor control

## Slide 5
Tank Level step response to 3-speed servo motor control
Zoomed into 3sec wait period
TP Uncertainty

## Slide 6: Where to measure resistance to determine “has switched” PASS/FAIL
LEVEL
Full
Empty
The Target Test Point Level (TP) will be set to ½ the level measurement uncertainty after 3 second settling time, prior to the specified upper bound of the Switch Zone.
REQUIREMENT: The Target Test Point uncertainty shall not exceed ???”.
To be defined

## Slide 7: Tank Level Response at various motor speeds
Using the Pololu MC18V7 PWM motor controller with Vs set to 24Vdc, The  Level was measure for 30 seconds at various percent of max speed (100% duty cycle) for both FILL and DRAIN at near EMPTY and near FULL.
The response has a minor shift at each case but is basically linear down to about 5% resulting in fill rate of about 0.1”/min where the control becomes unstable by 4% as shown in the ‘Zoomed In Plot.
Without doing an extensive fluid analysis, this response is assumed to be adequate for the MLOI application.
Data from mloi:\PM_Folder\_3 Develop\Design Calculations.xlsx sheet Tank. 3/4/2025 R. Ales

## Slide 8: Breadboard LT3092 Current Source Drift over 15 hours
One Plot (marker)

## Slide 9: Current source stability over 15 hrs, Start and End trace, 2000 seconds each trace
Time 14h26min – 15hrs ChA Blue, ChB Red
Mean 1.469V (14.78ma) 1.462V (14.66ma)
Time 0 -26min ChA Blue, ChB Red
Mean 1.468V (1.477ma) 1.463V (14.67ma)

## Slide 10
Motor
Controller
Pololu
18V7
Left
Right
DUT
Sensor
x4
DataQ
DI-2108
Buffer/
Scaler
x8
Laser
24Vdc
SMPS
24Vdc
Instrument
Power
Pump
4A DC Motor
LV Win11
Workstation
Interface PCB
Custom Unit
14ma
I-source
x8
Connector
USB Hub
TANK
NOT TO SCALE
To Laser & buffers
Keyance
LK-G2000
=
CONNECTOR EXAMPLE

## Slide 11
Motor
Controller
Pololu
18V7
Left
Right
DUT
Sensor
x4
DataQ
DI-2108
Buffer/
Scaler
x8
+5Vdc
USB Hub
Laser
Keyance
LG-K
24Vdc
SMPS
24Vdc
Instrument
Power
Pump
4A DC Motor
LV Win11
Workstation
Interface PCB
Custom Unit
14ma
I-source
x8
Connector

## Slide 12
-
+
-
+
Ain 0
Ain 1
Do-6
Do-5
Do-4
Tri-state
Buffer
Motor
Controller
On/Off
Fwd/Rev
Left
Right
DUT
Sensor
x4
DataQ
DI-2108
x4
+5Vdc

## Slide 13
To USB Hub
DATAQ
DI-2108
Interface
Circuit
Clip to
OLS DUT
To 24Vdc
R1
R2
R3
R4
R5
R6
R7
R8
R10
R9
U1A
U2
U1B
U3
1
3
2
1
3
2
1