## Slide 1
Step 2

## Slide 2
Step 5

## Slide 3
C:\Program Files (x86)\National Instruments\LabVIEW 2024\instr.lib\Keyence LK-G3000\Private\vi\Get DLL Path.vi

## Slide 4
Step 2

## Slide 5
1
2
3
SIM button

## Slide 6
Enter date here
Navigate to this Key
Edit this Value

## Slide 7
Click to START or END the automatic Test protocol
System State and Mode indicator
STOPS and Exit the application
GREEN Level
WHITE ChA
RED ChB

## Slide 8
With Simulate ON, the Level PV and Green Plot should simulate the tank level.
Set Point (SP) The desired Level in automatic.
Process Variable (PV) The measured Level.
With Pump JOG button:
ON, the FILL and DRAIN buttons and the Pump Speed dropdown control the pump.
OFF then the Level (PV) goes to the SP.
NOTE: SP, PV and Level indicate:
Tank is EMPTY when Level ~8.2”
Tank is FULL when Level ~ 2.4”

## Slide 9
ChA Value.
ChB Value.
Level Value.

## Slide 10
When Running (Green) The test protocol state machine sets the SP in automatic mode to create the stairstep level values.
FILL and DRAIN Status lights,
  GREEN = PASS
  RED = FAIL
  YELLOW = Test Error
  BLACK = Not Tested
ChA FILL
ChA DRAIN
ChB FILL
ChB DRAIN
Measured Resistance
(4 places)
Measured Target Test Level
(4 places)
Current Test State

## Slide 11
STOP and Exit button.
Test Information Dialog
Test Information Fields (7 Places)

## Slide 12
Config File Selection Field
Saved in Register
Currently Selected
Available Selection in current folder
Change Folder
Currently Selected
Full Path & Name

## Slide 13
Test Information

## Slide 14
WinDaq Data Acquisition Software
PicoScope oscilloscope software and PicoLog data logging software
Pololu Simple Motor Controller 18v7 (Fully Assembled)

## Slide 15
Switching between LibUSB and CDC Modes The native communication mode of the DI-2108-P is LibUSB but it can be switched to the USB CDC (Communication Device Class). CDC mode allows the USB port of DI-2108-P devices to appear like a traditional RS-232 port, which is common across most operating systems and development languages. Please note: LibUSB communication is required for WinDaq Acquisition software and programming with our .NET SDK. Only switch to CDC mode when writing your own programs at the protocol level (for example, Python applications). Use the following sequence to switch your device to/from LibUSB/CDC modes:
Connect the device to your PC via the USB port.
Apply power to the device.
When mode LED stops blinking white and is in Idle mode, push and hold the Control button. This must be completed within five seconds after the device is in Idle mode.
When the mode LED turns red, release the Control button. 5. The Mode LED will flash white then indicate Idle in either LibUSB or CDC mode: • CDC mode: Blinking Yellow • LibUSB mode: Blinking Green

## Slide 16: Fluke mA 3/28 to 4/1

## Slide 22
DOE SIPOC
DOE Design Matrix

## Slide 25
100
200
Spec 1 Fill Hi
LCL
Spec 1 Fill Lo
UCL
Switch Zone
Target TP
(Test point)
Target TP Uncertainty
Test Set Point
Switch Point
RNG_Stop
100
200
Spec 1 Drain Hi
Spec 1 Drain Lo
Switch Zone
Target TP
(Test point)
Target TP Uncertainty
Test Set Point
Switch
Point
RNG_Stop
LEVEL
Full
Empty
LEVEL
Full
Empty

## Slide 26
Numeric
Outputs
Test Protocol
Current State
Ch A & Ch B
Resistance
Level Filling Shown
(opposite of ATP Spec}
FULL
EMPTY

## Slide 27
Status
Mode Select
Manual
Controls
Automatic
Controls

## Slide 28
Raw Laser Value (per ATP)
Depth Gauge Reading
Depth Gauge Reading

## Slide 31: MLOI Gage R & R Evaluation Plan
B
C
3 Appraisers
A
+2σ
-2σ
Level Calibration
3 Appraisers
5 Replicates
Resistance Calibration
1 Appraiser
5 Replicates
+
+
Process/Part Variation
5 Sensors
Cal 1 Cal 2 Cal 3
Gage R & R
Step 1: Have 3 different Appraisers perform Calibration.
Step 2: Use the mean and 95% Limits for the Gage R&R for AV.
Step 3: One Appraiser performs 5 calibration of the resistance.
Step 4: Use the mean resistance calibration.
Step 5: Create two randomized test plans of 3 replicates each.
Step 6: Run Test Plan 1 in the AM, Run Test Plan 2 in the PM.
Total Measured Variation
3 Level Calibrations
5 Sensors
6 Replicates
1 Resistance Calibration
TOTAL 90 Measurements
=

## Slide 32: MLOI Gage R & R Evaluation Plan
Level Calibration
3 Appraisers
5 Replicates
Resistance Calibration
1 Appraiser
5 Replicates
Process/Part Variation
5 Sensors
Gage R & R
Step 1: Have 3 different Appraisers perform Calibration.
Step 2: Use the mean and 95% Limits for the Gage R&R for AV.
Step 3: One Appraiser performs 5 calibration of the resistance.
Step 4: Use the mean resistance calibration.
Step 5: Create two randomized test plans of 3 replicates each.
Step 6: Run Test Plan 1 in the AM, Run Test Plan 2 in the PM.
Total Measured Variation
3 Level Calibrations
5 Sensors
6 Replicates
1 Resistance Calibration
TOTAL 90 Measurements
=
EV = Sample_Range_Avg/d2(1.692)
AV = √ Cal_Range_Avg/ d2(1.912) –(EV)2/(n*r)
VGRR = √EV2 + AV2
PV = Part_Range_Avg//d2(3.1791)
Goodness % = VGRR/PV < 10%

## Slide 34
Voltage readings are highlighted yellow
Ohm readings are highlighted green

## Slide 35
Resistor
Calibration

## Slide 36
ChA PLUS
ChA GND
ChB PLUS
ChB
GND
All these bus wires are GND
ChA PLUS
ChB PLUS
R100-A
R100-B
R1K-A
R1K-B
ChA PLUS
ChB PLUS
GND