## Switch 1A
[11 empty] | Operators: | 3
[11 empty] | Trials: | 3
[11 empty] | Samples: | 10
Part No.: | 8005571.05 |  |  | Gage No.: |  |  |  |  | Date:
Part Name: | Oil Level Sensor |  |  | Gage Type: | Tank |  |  |  | Performed by:
Spec. | Min: | 98 |  | Nom: | 100 |  | Max: | 102 |  | Tolerance: | 4
| Operator A |  |  |  | Operator B |  |  |  | Operator C |  |  |  |  | R (barbar) | UCLR
Test Unit# | Trial #1 | Trial #2 | Trial #3 | Range | Trial #1 | Trial #2 | Trial #3 | Range | Trial #1 | Trial #2 | Trial #3 | Range |  | 0.24 | 0.6192
14 | 100 | 100.1 | 100.1 | 0.1 | 100 | 100 | 100 | 0 | 100.4 | 100 | 100 | 0.4 |  | 0.24 | 0.6192
15 | 100.2 | 100.3 | 100.3 | 0.1 | 100 | 100.3 | 100.3 | 0.3 | 100.7 | 100.3 | 100.3 | 0.4 |  | 0.24 | 0.6192
12 | 100.5 | 100.7 | 100.6 | 0.2 | 101 | 100.6 | 100.5 | 0.5 | 100.7 | 100.5 | 100.6 | 0.2 |  | 0.24 | 0.6192
11 | 100.5 | 100.6 | 100.6 | 0.1 | 101 | 100.6 | 100.6 | 0.4 | 100.9 | 100.6 | 100.6 | 0.3 |  | 0.24 | 0.6192
30 | 100.1 | 100.1 | 100.1 | 0 | 100 | 100.1 | 100.1 | 0.1 | 100.4 | 100.1 | 100.1 | 0.3 |  | 0.24 | 0.6192
31 | 100 | 100 | 100 | 0 | 100 | 100 | 100 | 0 | 100.3 | 100 | 100 | 0.3 |  | 0.24 | 0.6192
3 | 100.4 | 100.5 | 100.5 | 0.1 | 101 | 100.5 | 100.5 | 0.5 | 101.1 | 100.5 | 100.5 | 0.6 |  | 0.24 | 0.6192
2 | 100.7 | 100.7 | 100.7 | 0 | 101 | 100.7 | 100.7 | 0.3 | 101.3 | 100.7 | 100.7 | 0.6 |  | 0.24 | 0.6192
1 | 100.4 | 100.4 | 100.4 | 0 | 101 | 100.4 | 100.4 | 0.6 | 100.7 | 100.4 | 100.4 | 0.3 |  | 0.24 | 0.6192
29 | 100.1 | 100.1 | 100.1 | 0 | 100 | 100.1 | 100.1 | 0.1 | 100.5 | 100.1 | 100.1 | 0.4
Total | 1002.9 | 1003.5 | 1003.4 | 0.6 | 1005 | 1003.3 | 1003.2 | 2.8 | 1007 | 1003.2 | 1003.3 | 3.8
| X-bar (A): | 100.3267 | R-bar (A): | 0.06 | X-bar (B): | 100.3833 | R-bar (B): | 0.28 | X-bar (C) | 100.45 | R-bar (C) | 0.38
| R-bar-bar | 0.24 |  | UCLR | 0.6192 |  | X-bar (Diff.) | 0.12333
`
Repeatability - Equipment Variation (E.V.) |  |  |  |  | E.V. | 0.732 |  | SD(EV) | 0.14214 |  | % E.V. | 18.3%
Reproducibility - Appraiser Variation (A.V.) |  |  |  |  | A.V. | 0.30501 |  | SD(AV) | 0.059224 |  | % A.V. | 7.625%
Repeatability and Reproducibility (R&R) |  |  |  |  | R&R | 0.793 |  | SD(R&R) | 0.15398 |  | % R&R | 19.83%
[5 empty] | OVERALL GAGE R&R RESULTS: | GOOD
Repeatability and Reproducibility Study
(Gage R&R)

## Switch 2A
[11 empty] | Operators: | 3
[11 empty] | Trials: | 3
[11 empty] | Samples: | 10
Part No.: | 8005571.05 |  |  | Gage No.: |  |  |  |  | Date:
Part Name: | Oil Level Sensor |  |  | Gage Type: | Tank |  |  |  | Performed by:
Spec. | Min: | 197 |  | Nom: | 200 |  | Max: | 203 |  | Tolerance: | 6
| Operator A |  |  |  | Operator B |  |  |  | Operator C |  |  |  |  | R (barbar) | UCLR
Test Unit# | Trial #1 | Trial #2 | Trial #3 | Range | Trial #1 | Trial #2 | Trial #3 | Range | Trial #1 | Trial #2 | Trial #3 | Range |  | 0.20333 | 0.5246
14 | 200.1 | 200.1 | 200.1 | 0 | 200 | 200.1 | 200.1 | 0.1 | 200.4 | 200.1 | 200.1 | 0.3 |  | 0.20333 | 0.5246
15 | 200.1 | 200.2 | 200.1 | 0.1 | 200 | 200.1 | 200.1 | 0.1 | 200.5 | 200.1 | 200.1 | 0.4 |  | 0.20333 | 0.5246
12 | 200.8 | 201 | 201 | 0.2 | 201 | 200.9 | 200.9 | 0.1 | 201.1 | 200.9 | 200.9 | 0.2 |  | 0.20333 | 0.5246
11 | 200.9 | 201 | 201 | 0.1 | 201 | 201 | 201 | 0 | 201.3 | 201 | 201 | 0.3 |  | 0.20333 | 0.5246
30 | 200.4 | 200.2 | 200.3 | 0.2 | 200 | 200.2 | 200.2 | 0.2 | 200.7 | 200.4 | 200.5 | 0.3 |  | 0.20333 | 0.5246
31 | 199.8 | 200 | 200 | 0.2 | 200 | 199.9 | 199.9 | 0.1 | 200.2 | 199.9 | 199.9 | 0.3 |  | 0.20333 | 0.5246
3 | 200.8 | 200.9 | 200.9 | 0.1 | 201 | 200.8 | 200.9 | 0.2 | 201.4 | 200.8 | 200.9 | 0.6 |  | 0.20333 | 0.5246
2 | 201 | 201.1 | 201.1 | 0.1 | 201 | 201.1 | 201.1 | 0.1 | 201.7 | 201.1 | 201.1 | 0.6 |  | 0.20333 | 0.5246
1 | 200.7 | 200.7 | 200.7 | 0 | 201 | 200.7 | 200.7 | 0.3 | 201 | 200.7 | 200.7 | 0.3 |  | 0.20333 | 0.5246
29 | 200.2 | 200.2 | 200.2 | 0 | 200 | 200.2 | 200.2 | 0.2 | 200.6 | 200.2 | 200.2 | 0.4
Total | 2004.8 | 2005.4 | 2005.4 | 1 | 2005 | 2005 | 2005.1 | 1.4 | 2008.9 | 2005.2 | 2005.4 | 3.7
| X-bar (A): | 200.52 | R-bar (A): | 0.1 | X-bar (B): | 200.5033 | R-bar (B): | 0.14 | X-bar (C) | 200.65 | R-bar (C) | 0.37
| R-bar-bar | 0.20333 |  | UCLR | 0.5246 |  | X-bar (Diff.) | 0.14667
`
Repeatability - Equipment Variation (E.V.) |  |  |  |  | E.V. | 0.62017 |  | SD(EV) | 0.12042 |  | % E.V. | 10.34%
Reproducibility - Appraiser Variation (A.V.) |  |  |  |  | A.V. | 0.37947 |  | SD(AV) | 0.073683 |  | % A.V. | 6.324%
Repeatability and Reproducibility (R&R) |  |  |  |  | R&R | 0.72705 |  | SD(R&R) | 0.14117 |  | % R&R | 12.12%
[5 empty] | OVERALL GAGE R&R RESULTS: | GOOD
Repeatability and Reproducibility Study
(Gage R&R)

## Switch 3A
[11 empty] | Operators: | 3
[11 empty] | Trials: | 3
[11 empty] | Samples: | 10
Part No.: | 8005571.05 |  |  | Gage No.: |  |  |  |  | Date:
Part Name: | Oil Level Sensor |  |  | Gage Type: | Tank |  |  |  | Performed by:
Spec. | Min: | 296 |  | Nom: | 300 |  | Max: | 304 |  | Tolerance: | 8
| Operator A |  |  |  | Operator B |  |  |  | Operator C |  |  |  |  | R (barbar) | UCLR
Test Unit# | Trial #1 | Trial #2 | Trial #3 | Range | Trial #1 | Trial #2 | Trial #3 | Range | Trial #1 | Trial #2 | Trial #3 | Range |  | 0.22 | 0.5676
14 | 300 | 300 | 300 | 0 | 300 | 300 | 300 | 0 | 300.4 | 300 | 300 | 0.4 |  | 0.22 | 0.5676
15 | 300.2 | 300.3 | 300.2 | 0.1 | 300 | 300.2 | 300.2 | 0.2 | 300.6 | 300.2 | 300.2 | 0.4 |  | 0.22 | 0.5676
12 | 301 | 301.2 | 301.2 | 0.2 | 301 | 301.1 | 301.1 | 0.1 | 301.2 | 301.1 | 301.1 | 0.1 |  | 0.22 | 0.5676
11 | 301.2 | 301.4 | 301.4 | 0.2 | 301 | 301.3 | 301.3 | 0.3 | 301.6 | 301.3 | 301.4 | 0.3 |  | 0.22 | 0.5676
30 | 299.6 | 299.7 | 299.7 | 0.1 | 300 | 299.7 | 299.6 | 0.4 | 300 | 299.7 | 299.7 | 0.3 |  | 0.22 | 0.5676
31 | 299.9 | 300 | 300 | 0.1 | 300 | 299.9 | 299.9 | 0.1 | 300.2 | 299.9 | 299.9 | 0.3 |  | 0.22 | 0.5676
3 | 300.9 | 301.1 | 301 | 0.2 | 301 | 301 | 301 | 0 | 301.6 | 301.1 | 301 | 0.6 |  | 0.22 | 0.5676
2 | 301.2 | 301.3 | 301.3 | 0.1 | 301 | 301.3 | 301.3 | 0.3 | 302 | 301.3 | 301.3 | 0.7 |  | 0.22 | 0.5676
1 | 300.7 | 300.7 | 300.7 | 0 | 301 | 300.7 | 300.7 | 0.3 | 300.9 | 300.7 | 300.7 | 0.2 |  | 0.22 | 0.5676
29 | 300.1 | 300.2 | 300.1 | 0.1 | 300 | 300.2 | 300.2 | 0.2 | 300.5 | 300.2 | 300.2 | 0.3
Total | 3004.8 | 3005.9 | 3005.6 | 1.1 | 3005 | 3005.4 | 3005.3 | 1.9 | 3009 | 3005.5 | 3005.5 | 3.6
| X-bar (A): | 300.5433 | R-bar (A): | 0.11 | X-bar (B): | 300.5233 | R-bar (B): | 0.19 | X-bar (C) | 300.6667 | R-bar (C) | 0.36
| R-bar-bar | 0.22 |  | UCLR | 0.5676 |  | X-bar (Diff.) | 0.14333
`
Repeatability - Equipment Variation (E.V.) |  |  |  |  | E.V. | 0.671 |  | SD(EV) | 0.13029 |  | % E.V. | 8.388%
Reproducibility - Appraiser Variation (A.V.) |  |  |  |  | A.V. | 0.3671 |  | SD(AV) | 0.071281 |  | % A.V. | 4.589%
Repeatability and Reproducibility (R&R) |  |  |  |  | R&R | 0.76485 |  | SD(R&R) | 0.14852 |  | % R&R | 9.561%
[5 empty] | OVERALL GAGE R&R RESULTS: | EXCELLENT
Repeatability and Reproducibility Study
(Gage R&R)

## Switch 4A
[11 empty] | Operators: | 3
[11 empty] | Trials: | 3
[11 empty] | Samples: | 10
Part No.: | 8005571.05 |  |  | Gage No.: |  |  |  |  | Date:
Part Name: | Oil Level Sensor |  |  | Gage Type: | Tank |  |  |  | Performed by:
Spec. | Min: | 395 |  | Nom: | 400 |  | Max: | 405 |  | Tolerance: | 10
| Operator A |  |  |  | Operator B |  |  |  | Operator C |  |  |  |  | R (barbar) | UCLR
Test Unit# | Trial #1 | Trial #2 | Trial #3 | Range | Trial #1 | Trial #2 | Trial #3 | Range | Trial #1 | Trial #2 | Trial #3 | Range |  | 0.25333 | 0.6536
14 | 399.8 | 399.9 | 399.9 | 0.1 | 400 | 399.8 | 399.8 | 0.2 | 400.2 | 399.8 | 399.8 | 0.4 |  | 0.25333 | 0.6536
15 | 400 | 400.2 | 400.1 | 0.2 | 400 | 400.1 | 400 | 0.1 | 400.4 | 400 | 400 | 0.4 |  | 0.25333 | 0.6536
12 | 401.3 | 401.5 | 401.5 | 0.2 | 401 | 401.4 | 401.4 | 0.4 | 401.5 | 401.4 | 401.4 | 0.1 |  | 0.25333 | 0.6536
11 | 401.5 | 401.6 | 401.7 | 0.2 | 402 | 401.6 | 401.6 | 0.4 | 401.9 | 401.6 | 401.6 | 0.3 |  | 0.25333 | 0.6536
30 | 399.5 | 399.5 | 399.5 | 0 | 400 | 399.5 | 399.5 | 0.5 | 399.8 | 399.5 | 399.5 | 0.3 |  | 0.25333 | 0.6536
31 | 399.7 | 399.8 | 399.8 | 0.1 | 400 | 399.7 | 399.8 | 0.3 | 400.1 | 399.8 | 399.8 | 0.3 |  | 0.25333 | 0.6536
3 | 401.2 | 401.3 | 401.3 | 0.1 | 401 | 401.3 | 401.3 | 0.3 | 401.9 | 401.3 | 401.3 | 0.6 |  | 0.25333 | 0.6536
2 | 401.5 | 401.6 | 401.6 | 0.1 | 402 | 401.6 | 401.6 | 0.4 | 402.3 | 401.6 | 401.7 | 0.7 |  | 0.25333 | 0.6536
1 | 400.9 | 400.9 | 400.9 | 0 | 401 | 400.9 | 400.9 | 0.1 | 401.1 | 400.9 | 400.9 | 0.2 |  | 0.25333 | 0.6536
29 | 399.8 | 399.8 | 399.8 | 0 | 400 | 399.8 | 399.8 | 0.2 | 400.2 | 399.8 | 399.8 | 0.4
Total | 4005.2 | 4006.1 | 4006.1 | 1 | 4007 | 4005.7 | 4005.7 | 2.9 | 4009.4 | 4005.7 | 4005.8 | 3.7
| X-bar (A): | 400.58 | R-bar (A): | 0.1 | X-bar (B): | 400.6133 | R-bar (B): | 0.29 | X-bar (C) | 400.6967 | R-bar (C) | 0.37
| R-bar-bar | 0.25333 |  | UCLR | 0.6536 |  | X-bar (Diff.) | 0.11667
`
Repeatability - Equipment Variation (E.V.) |  |  |  |  | E.V. | 0.77267 |  | SD(EV) | 0.15003 |  | % E.V. | 7.727%
Reproducibility - Appraiser Variation (A.V.) |  |  |  |  | A.V. | 0.28165 |  | SD(AV) | 0.054689 |  | % A.V. | 2.816%
Repeatability and Reproducibility (R&R) |  |  |  |  | R&R | 0.8224 |  | SD(R&R) | 0.15969 |  | % R&R | 8.224%
[5 empty] | OVERALL GAGE R&R RESULTS: | EXCELLENT
Repeatability and Reproducibility Study
(Gage R&R)

## Switch 5A
[11 empty] | Operators: | 3
[11 empty] | Trials: | 3
[11 empty] | Samples: | 10
Part No.: | 8005571.05 |  |  | Gage No.: |  |  |  |  | Date:
Part Name: | Oil Level Sensor |  |  | Gage Type: | Tank |  |  |  | Performed by:
Spec. | Min: | 494 |  | Nom: | 500 |  | Max: | 506 |  | Tolerance: | 12
| Operator A |  |  |  | Operator B |  |  |  | Operator C |  |  |  |  | R (barbar) | UCLR
Test Unit# | Trial #1 | Trial #2 | Trial #3 | Range | Trial #1 | Trial #2 | Trial #3 | Range | Trial #1 | Trial #2 | Trial #3 | Range |  | 0.22333 | 0.5762
14 | 499.8 | 499.8 | 499.8 | 0 | 500 | 499.8 | 499.8 | 0.2 | 500.1 | 499.8 | 499.8 | 0.3 |  | 0.22333 | 0.5762
15 | 500.1 | 500.2 | 500.1 | 0.1 | 500 | 500.1 | 500.1 | 0.1 | 500.5 | 500.1 | 500.1 | 0.4 |  | 0.22333 | 0.5762
12 | 501.9 | 502.2 | 502.2 | 0.3 | 502 | 502 | 502.1 | 0.1 | 502.2 | 502.1 | 502.1 | 0.1 |  | 0.22333 | 0.5762
11 | 501.6 | 501.8 | 501.8 | 0.2 | 502 | 501.8 | 501.8 | 0.2 | 502 | 501.8 | 501.8 | 0.2 |  | 0.22333 | 0.5762
30 | 499.2 | 499.3 | 499.3 | 0.1 | 499 | 499.2 | 499.2 | 0.2 | 499.6 | 499.3 | 499.3 | 0.3 |  | 0.22333 | 0.5762
31 | 499.7 | 499.7 | 499.7 | 0 | 500 | 499.7 | 499.7 | 0.3 | 500 | 499.7 | 499.7 | 0.3 |  | 0.22333 | 0.5762
3 | 501.4 | 501.6 | 501.6 | 0.2 | 502 | 501.5 | 501.6 | 0.5 | 502.1 | 501.6 | 501.6 | 0.5 |  | 0.22333 | 0.5762
2 | 501.8 | 502 | 502 | 0.2 | 502 | 502 | 502 | 0 | 502.6 | 502 | 502 | 0.6 |  | 0.22333 | 0.5762
1 | 501.1 | 501.1 | 501.1 | 0 | 501 | 501.1 | 501.1 | 0.1 | 501.3 | 501.1 | 501.1 | 0.2 |  | 0.22333 | 0.5762
29 | 499.5 | 499.5 | 499.6 | 0.1 | 500 | 499.5 | 499.5 | 0.5 | 499.9 | 499.5 | 499.6 | 0.4
Total | 5006.1 | 5007.2 | 5007.2 | 1.2 | 5008 | 5006.7 | 5006.9 | 2.2 | 5010.3 | 5007 | 5007.1 | 3.3
| X-bar (A): | 500.6833 | R-bar (A): | 0.12 | X-bar (B): | 500.72 | R-bar (B): | 0.22 | X-bar (C) | 500.8133 | R-bar (C) | 0.33
| R-bar-bar | 0.22333 |  | UCLR | 0.5762 |  | X-bar (Diff.) | 0.13
`
Repeatability - Equipment Variation (E.V.) |  |  |  |  | E.V. | 0.68117 |  | SD(EV) | 0.13227 |  | % E.V. | 5.676%
Reproducibility - Appraiser Variation (A.V.) |  |  |  |  | A.V. | 0.32823 |  | SD(AV) | 0.063734 |  | % A.V. | 2.735%
Repeatability and Reproducibility (R&R) |  |  |  |  | R&R | 0.75612 |  | SD(R&R) | 0.14682 |  | % R&R | 6.301%
[5 empty] | OVERALL GAGE R&R RESULTS: | EXCELLENT
Repeatability and Reproducibility Study
(Gage R&R)

## Switch 6A
[11 empty] | Operators: | 3
[11 empty] | Trials: | 3
[11 empty] | Samples: | 10
Part No.: | 8005571.05 |  |  | Gage No.: |  |  |  |  | Date:
Part Name: | Oil Level Sensor |  |  | Gage Type: | Tank |  |  |  | Performed by:
Spec. | Min: | 594 |  | Nom: | 600 |  | Max: | 606 |  | Tolerance: | 12
| Operator A |  |  |  | Operator B |  |  |  | Operator C |  |  |  |  | R (barbar) | UCLR
Test Unit# | Trial #1 | Trial #2 | Trial #3 | Range | Trial #1 | Trial #2 | Trial #3 | Range | Trial #1 | Trial #2 | Trial #3 | Range |  | 0.2 | 0.516
14 | 600 | 600 | 600 | 0 | 600 | 600 | 600 | 0 | 600 | 600 | 600 | 0 |  | 0.2 | 0.516
15 | 600 | 600 | 600 | 0 | 600 | 600 | 600 | 0 | 600 | 600 | 600 | 0 |  | 0.2 | 0.516
12 | 602 | 602 | 603 | 1 | 602 | 602 | 602 | 0 | 603 | 602 | 602 | 1 |  | 0.2 | 0.516
11 | 602 | 602 | 602 | 0 | 602 | 602 | 602 | 0 | 602 | 602 | 602 | 0 |  | 0.2 | 0.516
30 | 599 | 599 | 599 | 0 | 599 | 599 | 599 | 0 | 599 | 599 | 599 | 0 |  | 0.2 | 0.516
31 | 600 | 600 | 600 | 0 | 600 | 600 | 600 | 0 | 600 | 600 | 600 | 0 |  | 0.2 | 0.516
3 | 602 | 602 | 602 | 0 | 602 | 602 | 602 | 0 | 602 | 602 | 602 | 0 |  | 0.2 | 0.516
2 | 602 | 602 | 602 | 0 | 602 | 602 | 603 | 1 | 603 | 602 | 602 | 1 |  | 0.2 | 0.516
1 | 601 | 601 | 601 | 0 | 601 | 601 | 601 | 0 | 602 | 601 | 601 | 1 |  | 0.2 | 0.516
29 | 599 | 599 | 599 | 0 | 599 | 599 | 599 | 0 | 600 | 599 | 599 | 1
Total | 6007 | 6007 | 6008 | 1 | 6007 | 6007 | 6008 | 1 | 6011 | 6007 | 6007 | 4
| X-bar (A): | 600.7333 | R-bar (A): | 0.1 | X-bar (B): | 600.7333 | R-bar (B): | 0.1 | X-bar (C) | 600.8333 | R-bar (C) | 0.4
| R-bar-bar | 0.2 |  | UCLR | 0.516 |  | X-bar (Diff.) | 0.1
`
Repeatability - Equipment Variation (E.V.) |  |  |  |  | E.V. | 0.61 |  | SD(EV) | 0.11845 |  | % E.V. | 5.083%
Reproducibility - Appraiser Variation (A.V.) |  |  |  |  | A.V. | 0.24596 |  | SD(AV) | 0.047759 |  | % A.V. | 2.05%
Repeatability and Reproducibility (R&R) |  |  |  |  | R&R | 0.65772 |  | SD(R&R) | 0.12771 |  | % R&R | 5.481%
[5 empty] | OVERALL GAGE R&R RESULTS: | EXCELLENT
Repeatability and Reproducibility Study
(Gage R&R)

## Switch 7A
[11 empty] | Operators: | 3
[11 empty] | Trials: | 3
[11 empty] | Samples: | 10
Part No.: | 8005571.05 |  |  | Gage No.: |  |  |  |  | Date:
Part Name: | Oil Level Sensor |  |  | Gage Type: | Tank |  |  |  | Performed by:
Spec. | Min: | 692 |  | Nom: | 700 |  | Max: | 708 |  | Tolerance: | 16
| Operator A |  |  |  | Operator B |  |  |  | Operator C |  |  |  |  | R (barbar) | UCLR
Test Unit# | Trial #1 | Trial #2 | Trial #3 | Range | Trial #1 | Trial #2 | Trial #3 | Range | Trial #1 | Trial #2 | Trial #3 | Range |  | 0.23333 | 0.602
14 | 700 | 700 | 700 | 0 | 700 | 700 | 700 | 0 | 700 | 700 | 700 | 0 |  | 0.23333 | 0.602
15 | 700 | 700 | 700 | 0 | 700 | 700 | 700 | 0 | 700 | 700 | 700 | 0 |  | 0.23333 | 0.602
12 | 703 | 703 | 703 | 0 | 703 | 703 | 703 | 0 | 703 | 703 | 703 | 0 |  | 0.23333 | 0.602
11 | 702 | 702 | 702 | 0 | 702 | 703 | 702 | 1 | 703 | 702 | 702 | 1 |  | 0.23333 | 0.602
30 | 699 | 699 | 699 | 0 | 699 | 699 | 699 | 0 | 699 | 699 | 699 | 0 |  | 0.23333 | 0.602
31 | 700 | 700 | 700 | 0 | 700 | 699 | 700 | 1 | 700 | 700 | 700 | 0 |  | 0.23333 | 0.602
3 | 702 | 702 | 702 | 0 | 702 | 702 | 702 | 0 | 703 | 702 | 702 | 1 |  | 0.23333 | 0.602
2 | 702 | 703 | 703 | 1 | 703 | 703 | 703 | 0 | 703 | 703 | 703 | 0 |  | 0.23333 | 0.602
1 | 701 | 701 | 701 | 0 | 701 | 701 | 701 | 0 | 702 | 701 | 701 | 1 |  | 0.23333 | 0.602
29 | 699 | 699 | 699 | 0 | 699 | 699 | 699 | 0 | 700 | 699 | 699 | 1
Total | 7008 | 7009 | 7009 | 1 | 7009 | 7009 | 7009 | 2 | 7013 | 7009 | 7009 | 4
| X-bar (A): | 700.8667 | R-bar (A): | 0.1 | X-bar (B): | 700.9 | R-bar (B): | 0.2 | X-bar (C) | 701.0333 | R-bar (C) | 0.4
| R-bar-bar | 0.23333 |  | UCLR | 0.602 |  | X-bar (Diff.) | 0.16667
`
Repeatability - Equipment Variation (E.V.) |  |  |  |  | E.V. | 0.71167 |  | SD(EV) | 0.13819 |  | % E.V. | 4.448%
Reproducibility - Appraiser Variation (A.V.) |  |  |  |  | A.V. | 0.43083 |  | SD(AV) | 0.083657 |  | % A.V. | 2.693%
Repeatability and Reproducibility (R&R) |  |  |  |  | R&R | 0.83192 |  | SD(R&R) | 0.16154 |  | % R&R | 5.199%
[5 empty] | OVERALL GAGE R&R RESULTS: | EXCELLENT
Repeatability and Reproducibility Study
(Gage R&R)

## Switch 8A
[11 empty] | Operators: | 3
[11 empty] | Trials: | 3
[11 empty] | Samples: | 10
Part No.: | 8005571.05 |  |  | Gage No.: |  |  |  |  | Date:
Part Name: | Oil Level Sensor |  |  | Gage Type: | Tank |  |  |  | Performed by:
Spec. | Min: | 791 |  | Nom: | 800 |  | Max: | 809 |  | Tolerance: | 18
| Operator A |  |  |  | Operator B |  |  |  | Operator C |  |  |  |  | R (barbar) | UCLR
Test Unit# | Trial #1 | Trial #2 | Trial #3 | Range | Trial #1 | Trial #2 | Trial #3 | Range | Trial #1 | Trial #2 | Trial #3 | Range |  | 0.13333 | 0.344
14 | 800 | 800 | 800 | 0 | 800 | 800 | 800 | 0 | 800 | 800 | 799 | 1 |  | 0.13333 | 0.344
15 | 800 | 800 | 800 | 0 | 800 | 800 | 800 | 0 | 800 | 800 | 800 | 0 |  | 0.13333 | 0.344
12 | 803 | 803 | 803 | 0 | 803 | 803 | 803 | 0 | 803 | 803 | 803 | 0 |  | 0.13333 | 0.344
11 | 803 | 803 | 803 | 0 | 803 | 803 | 803 | 0 | 803 | 803 | 803 | 0 |  | 0.13333 | 0.344
30 | 799 | 799 | 799 | 0 | 799 | 799 | 799 | 0 | 799 | 799 | 798 | 1 |  | 0.13333 | 0.344
31 | 800 | 800 | 800 | 0 | 800 | 800 | 800 | 0 | 800 | 800 | 800 | 0 |  | 0.13333 | 0.344
3 | 803 | 803 | 803 | 0 | 803 | 803 | 803 | 0 | 803 | 803 | 803 | 0 |  | 0.13333 | 0.344
2 | 803 | 803 | 803 | 0 | 803 | 803 | 803 | 0 | 804 | 803 | 803 | 1 |  | 0.13333 | 0.344
1 | 802 | 802 | 802 | 0 | 802 | 802 | 802 | 0 | 802 | 802 | 802 | 0 |  | 0.13333 | 0.344
29 | 799 | 799 | 799 | 0 | 799 | 799 | 799 | 0 | 800 | 799 | 799 | 1
Total | 8012 | 8012 | 8012 | 0 | 8012 | 8012 | 8012 | 0 | 8014 | 8012 | 8010 | 4
| X-bar (A): | 801.2 | R-bar (A): | 0 | X-bar (B): | 801.2 | R-bar (B): | 0 | X-bar (C) | 801.2 | R-bar (C) | 0.4
| R-bar-bar | 0.13333 |  | UCLR | 0.344 |  | X-bar (Diff.) | 0
`
Repeatability - Equipment Variation (E.V.) |  |  |  |  | E.V. | 0.40667 |  | SD(EV) | 0.078964 |  | % E.V. | 2.259%
Reproducibility - Appraiser Variation (A.V.) |  |  |  |  | A.V. | #NUM! |  | SD(AV) | #NUM! |  | % A.V. | #NUM!
Repeatability and Reproducibility (R&R) |  |  |  |  | R&R | #NUM! |  | SD(R&R) | #NUM! |  | % R&R | #NUM!
[5 empty] | OVERALL GAGE R&R RESULTS: | #NUM!
Repeatability and Reproducibility Study
(Gage R&R)

## Switch 9A
[11 empty] | Operators: | 3
[11 empty] | Trials: | 3
[11 empty] | Samples: | 10
Part No.: | 8005571.05 |  |  | Gage No.: |  |  |  |  | Date:
Part Name: | Oil Level Sensor |  |  | Gage Type: | Tank |  |  |  | Performed by:
Spec. | Min: | 890 |  | Nom: | 900 |  | Max: | 910 |  | Tolerance: | 20
| Operator A |  |  |  | Operator B |  |  |  | Operator C |  |  |  |  | R (barbar) | UCLR
Test Unit# | Trial #1 | Trial #2 | Trial #3 | Range | Trial #1 | Trial #2 | Trial #3 | Range | Trial #1 | Trial #2 | Trial #3 | Range |  | 0.26667 | 0.688
14 | 900 | 900 | 900 | 0 | 900 | 900 | 900 | 0 | 900 | 900 | 899 | 1 |  | 0.26667 | 0.688
15 | 900 | 900 | 900 | 0 | 900 | 900 | 900 | 0 | 900 | 900 | 899 | 1 |  | 0.26667 | 0.688
12 | 904 | 904 | 904 | 0 | 904 | 904 | 904 | 0 | 905 | 904 | 904 | 1 |  | 0.26667 | 0.688
11 | 903 | 903 | 903 | 0 | 903 | 903 | 903 | 0 | 904 | 903 | 903 | 1 |  | 0.26667 | 0.688
30 | 898 | 898 | 898 | 0 | 898 | 898 | 898 | 0 | 899 | 898 | 898 | 1 |  | 0.26667 | 0.688
31 | 900 | 900 | 900 | 0 | 900 | 900 | 900 | 0 | 900 | 900 | 900 | 0 |  | 0.26667 | 0.688
3 | 903 | 903 | 903 | 0 | 903 | 903 | 903 | 0 | 904 | 903 | 903 | 1 |  | 0.26667 | 0.688
2 | 903 | 903 | 903 | 0 | 903 | 904 | 903 | 1 | 904 | 903 | 903 | 1 |  | 0.26667 | 0.688
1 | 902 | 902 | 902 | 0 | 902 | 902 | 902 | 0 | 902 | 902 | 902 | 0 |  | 0.26667 | 0.688
29 | 899 | 899 | 899 | 0 | 899 | 899 | 899 | 0 | 899 | 899 | 899 | 0
Total | 9012 | 9012 | 9012 | 0 | 9012 | 9013 | 9012 | 1 | 9017 | 9012 | 9010 | 7
| X-bar (A): | 901.2 | R-bar (A): | 0 | X-bar (B): | 901.2333 | R-bar (B): | 0.1 | X-bar (C) | 901.3 | R-bar (C) | 0.7
| R-bar-bar | 0.26667 |  | UCLR | 0.688 |  | X-bar (Diff.) | 0.1
`
Repeatability - Equipment Variation (E.V.) |  |  |  |  | E.V. | 0.81333 |  | SD(EV) | 0.15793 |  | % E.V. | 4.067%
Reproducibility - Appraiser Variation (A.V.) |  |  |  |  | A.V. | 0.2255 |  | SD(AV) | 0.043786 |  | % A.V. | 1.127%
Repeatability and Reproducibility (R&R) |  |  |  |  | R&R | 0.84401 |  | SD(R&R) | 0.16389 |  | % R&R | 4.22%
[5 empty] | OVERALL GAGE R&R RESULTS: | EXCELLENT
Repeatability and Reproducibility Study
(Gage R&R)

## Switch 10A
[11 empty] | Operators: | 3
[11 empty] | Trials: | 3
[11 empty] | Samples: | 10
Part No.: | 8005571.05 |  |  | Gage No.: |  |  |  |  | Date:
Part Name: | Oil Level Sensor |  |  | Gage Type: | Tank |  |  |  | Performed by:
Spec. | Min: | 989 |  | Nom: | 1000 |  | Max: | 1011 |  | Tolerance: | 22
| Operator A |  |  |  | Operator B |  |  |  | Operator C |  |  |  |  | R (barbar) | UCLR
Test Unit# | Trial #1 | Trial #2 | Trial #3 | Range | Trial #1 | Trial #2 | Trial #3 | Range | Trial #1 | Trial #2 | Trial #3 | Range |  | 0.36667 | 0.946
14 | 999 | 999 | 999 | 0 | 999 | 999 | 999 | 0 | 1000 | 999 | 999 | 1 |  | 0.36667 | 0.946
15 | 1000 | 1000 | 1000 | 0 | 1000 | 1000 | 1000 | 0 | 1000 | 1000 | 999 | 1 |  | 0.36667 | 0.946
12 | 1004 | 1005 | 1005 | 1 | 1004 | 1004 | 1005 | 1 | 1005 | 1005 | 1004 | 1 |  | 0.36667 | 0.946
11 | 1004 | 1004 | 1004 | 0 | 1004 | 1004 | 1004 | 0 | 1004 | 1004 | 1004 | 0 |  | 0.36667 | 0.946
30 | 998 | 998 | 998 | 0 | 998 | 998 | 998 | 0 | 999 | 998 | 998 | 1 |  | 0.36667 | 0.946
31 | 1000 | 1000 | 1000 | 0 | 1000 | 999 | 999 | 1 | 1000 | 999 | 1000 | 1 |  | 0.36667 | 0.946
3 | 1003 | 1003 | 1003 | 0 | 1003 | 1003 | 1003 | 0 | 1004 | 1003 | 1003 | 1 |  | 0.36667 | 0.946
2 | 1004 | 1004 | 1004 | 0 | 1004 | 1004 | 1004 | 0 | 1005 | 1004 | 1004 | 1 |  | 0.36667 | 0.946
1 | 1002 | 1002 | 1002 | 0 | 1002 | 1002 | 1002 | 0 | 1002 | 1002 | 1002 | 0 |  | 0.36667 | 0.946
29 | 999 | 999 | 999 | 0 | 999 | 999 | 999 | 0 | 1000 | 999 | 999 | 1
Total | 10013 | 10014 | 10014 | 1 | 10013 | 10012 | 10013 | 2 | 10019 | 10013 | 10012 | 8
| X-bar (A): | 1001.3667 | R-bar (A): | 0.1 | X-bar (B): | 1001.2667 | R-bar (B): | 0.2 | X-bar (C) | 1001.4667 | R-bar (C) | 0.8
| R-bar-bar | 0.36667 |  | UCLR | 0.946 |  | X-bar (Diff.) | 0.2
`
Repeatability - Equipment Variation (E.V.) |  |  |  |  | E.V. | 1.1183 |  | SD(EV) | 0.21715 |  | % E.V. | 5.083%
Reproducibility - Appraiser Variation (A.V.) |  |  |  |  | A.V. | 0.49991 |  | SD(AV) | 0.09707 |  | % A.V. | 2.272%
Repeatability and Reproducibility (R&R) |  |  |  |  | R&R | 1.225 |  | SD(R&R) | 0.23786 |  | % R&R | 5.568%
[5 empty] | OVERALL GAGE R&R RESULTS: | EXCELLENT
Repeatability and Reproducibility Study
(Gage R&R)

## Switch 11A
[11 empty] | Operators: | 3
[11 empty] | Trials: | 3
[11 empty] | Samples: | 10
Part No.: | 8005571.05 |  |  | Gage No.: |  |  |  |  | Date:
Part Name: | Oil Level Sensor |  |  | Gage Type: | Tank |  |  |  | Performed by:
Spec. | Min: |  |  | Nom: |  |  | Max: |  |  | Tolerance: | 0
| Operator A |  |  |  | Operator B |  |  |  | Operator C |  |  |  |  | R (barbar) | UCLR
Test Unit# | Trial #1 | Trial #2 | Trial #3 | Range | Trial #1 | Trial #2 | Trial #3 | Range | Trial #1 | Trial #2 | Trial #3 | Range |  | 0 | 0
14 |  |  |  | 0 |  |  |  | 0 |  |  |  | 0 |  | 0 | 0
15 |  |  |  | 0 |  |  |  | 0 |  |  |  | 0 |  | 0 | 0
12 |  |  |  | 0 |  |  |  | 0 |  |  |  | 0 |  | 0 | 0
11 |  |  |  | 0 |  |  |  | 0 |  |  |  | 0 |  | 0 | 0
30 |  |  |  | 0 |  |  |  | 0 |  |  |  | 0 |  | 0 | 0
31 |  |  |  | 0 |  |  |  | 0 |  |  |  | 0 |  | 0 | 0
3 |  |  |  | 0 |  |  |  | 0 |  |  |  | 0 |  | 0 | 0
2 |  |  |  | 0 |  |  |  | 0 |  |  |  | 0 |  | 0 | 0
1 |  |  |  | 0 |  |  |  | 0 |  |  |  | 0 |  | 0 | 0
29 |  |  |  | 0 |  |  |  | 0 |  |  |  | 0
Total | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| X-bar (A): | 0 | R-bar (A): | 0 | X-bar (B): | 0 | R-bar (B): | 0 | X-bar (C) | 0 | R-bar (C) | 0
| R-bar-bar | 0 |  | UCLR | 0 |  | X-bar (Diff.) | 0
`
Repeatability - Equipment Variation (E.V.) |  |  |  |  | E.V. | 0 |  | SD(EV) | 0 |  | % E.V. | #DIV/0!
Reproducibility - Appraiser Variation (A.V.) |  |  |  |  | A.V. | 0 |  | SD(AV) | 0 |  | % A.V. | #DIV/0!
Repeatability and Reproducibility (R&R) |  |  |  |  | R&R | 0 |  | SD(R&R) | 0 |  | % R&R | #DIV/0!
[5 empty] | OVERALL GAGE R&R RESULTS: | #DIV/0!
Repeatability and Reproducibility Study
(Gage R&R)

## Switch 12A
[11 empty] | Operators: | 3
[11 empty] | Trials: | 3
[11 empty] | Samples: | 10
Part No.: | 8005571.05 |  |  | Gage No.: |  |  |  |  | Date:
Part Name: | Oil Level Sensor |  |  | Gage Type: | Tank |  |  |  | Performed by:
Spec. | Min: |  |  | Nom: |  |  | Max: |  |  | Tolerance: | 0
| Operator A |  |  |  | Operator B |  |  |  | Operator C |  |  |  |  | R (barbar) | UCLR
Test Unit# | Trial #1 | Trial #2 | Trial #3 | Range | Trial #1 | Trial #2 | Trial #3 | Range | Trial #1 | Trial #2 | Trial #3 | Range |  | 0 | 0
14 |  |  |  | 0 |  |  |  | 0 |  |  |  | 0 |  | 0 | 0
15 |  |  |  | 0 |  |  |  | 0 |  |  |  | 0 |  | 0 | 0
12 |  |  |  | 0 |  |  |  | 0 |  |  |  | 0 |  | 0 | 0
11 |  |  |  | 0 |  |  |  | 0 |  |  |  | 0 |  | 0 | 0
30 |  |  |  | 0 |  |  |  | 0 |  |  |  | 0 |  | 0 | 0
31 |  |  |  | 0 |  |  |  | 0 |  |  |  | 0 |  | 0 | 0
3 |  |  |  | 0 |  |  |  | 0 |  |  |  | 0 |  | 0 | 0
2 |  |  |  | 0 |  |  |  | 0 |  |  |  | 0 |  | 0 | 0
1 |  |  |  | 0 |  |  |  | 0 |  |  |  | 0 |  | 0 | 0
29 |  |  |  | 0 |  |  |  | 0 |  |  |  | 0
Total | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| X-bar (A): | 0 | R-bar (A): | 0 | X-bar (B): | 0 | R-bar (B): | 0 | X-bar (C) | 0 | R-bar (C) | 0
| R-bar-bar | 0 |  | UCLR | 0 |  | X-bar (Diff.) | 0
`
Repeatability - Equipment Variation (E.V.) |  |  |  |  | E.V. | 0 |  | SD(EV) | 0 |  | % E.V. | #DIV/0!
Reproducibility - Appraiser Variation (A.V.) |  |  |  |  | A.V. | 0 |  | SD(AV) | 0 |  | % A.V. | #DIV/0!
Repeatability and Reproducibility (R&R) |  |  |  |  | R&R | 0 |  | SD(R&R) | 0 |  | % R&R | #DIV/0!
[5 empty] | OVERALL GAGE R&R RESULTS: | #DIV/0!
Repeatability and Reproducibility Study
(Gage R&R)