## Raw Data
Dragan | 2024-11-06 13:17
Index | Set Point | Laser | Measure |  | Slope | Offset
1 | 1.1 | 1.102 | 8.495 |  | -0.996 | 9.591
2 | 2.7 | 2.664 | 6.938
3 | 4.3 | 4.26 | 5.347
4 | 5.8 | 5.756 | 3.857
5 | 7.2 | 7.249 | 2.374
George | 2024-11-06 13:27
Index | Set Point | Laser | Measure |  | Slope | Offset
1 | 1.1 | 1.102 | 8.499 |  | -0.996 | 9.595
2 | 2.7 | 2.735 | 6.876
3 | 4.2 | 4.188 | 5.417
4 | 5.8 | 5.793 | 3.822
5 | 7.2 | 7.248 | 2.384
Dragan | 2024-11-06 13:40
Index | Set Point | Laser | Measure |  | Slope | Offset
1 | 1.1 | 1.102 | 8.5 |  | -0.995 | 9.601
2 | 2.7 | 2.722 | 6.897
3 | 4.2 | 4.213 | 5.406
4 | 5.7 | 5.738 | 3.888
5 | 7.2 | 7.249 | 2.384
George | 2024-11-06 13:49
Index | Set Point | Laser | Measure |  | Slope | Offset
1 | 1.1 | 1.102 | 8.503 |  | -0.996 | 9.601
2 | 2.7 | 2.703 | 6.909
3 | 4.2 | 4.176 | 5.444
4 | 5.7 | 5.745 | 3.881
5 | 7.2 | 7.248 | 2.382
Dragan | 2024-11-06 13:59
Index | Set Point | Laser | Measure |  | Slope | Offset
1 | 1.1 | 1.102 | 8.504 |  | -0.996 | 9.602
2 | 2.7 | 2.676 | 6.936
3 | 4.2 | 4.178 | 5.442
4 | 5.7 | 5.715 | 3.908
5 | 7.2 | 7.249 | 2.381
8.504 | 2024-11-06 14:09
Index | Set Point | Laser | Measure |  | Slope | Offset
1 | 1.1 | 1.101 | 8.504 |  | -0.996 | 9.602
2 | 2.7 | 2.707 | 6.907
3 | 4.2 | 4.183 | 5.439
4 | 5.7 | 5.716 | 3.906
5 | 7.2 | 7.249 | 2.383
rick | 2024-11-17 09:07
Index | Set Point | Laser | Measure |  | Slope | Offset
1 | 1 | 1.002 | 0 |  | 0 | 0
2 | 2.5 | 2.53 | 0
3 | 4.1 | 4.092 | 0
4 | 5.5 | 5.502 | 0
5 | 7 | 7.002 | 0
george | 45613.6514
Index | Set Point | Laser | Measure |  | Slope | Offset
1 | 1.1 | 1.102 | 0 |  | 0 | 0
2 | 2.7 | 2.66 | 0
3 | 4.2 | 4.248 | 0
4 | 5.7 | 5.7 | 0
5 | 7.2 | 7.202 | 0
rob | 2024-11-17 15:56
Index | Set Point | Laser | Measure |  | Slope | Offset
1 | 1.1 | 1.102 | 0 |  | 0 | 0
2 | 2.723 | 2.724 | 0
3 | 4.175 | 4.174 | 0
4 | 5.683 | 5.682 | 0
5 | 7.2 | 7.204 | 0
2024-11-17 16:29
Index | Set Point | Laser | Measure |  | Slope | Offset
1 | 1.1 | 1.102 | 8 |  | -1.003 | 9.066
2 | 2.67 | 2.67 | 6.5
3 | 4.165 | 4.166 | 4.7
4 | 5.725 | 5.728 | 3.2
5 | 7.2 | 7.202 | 2
rick | 2024-11-27 15:07
Index | Set Point | Laser | Measure |  | Slope | Offset | Constant
1 | 1.1 | 1.101 | 8.5 |  | -1.03 | 9.272 | -1.03
2 | 2.696 | 2.695 | 6.2
3 | 4.166 | 4.165 | 4.8
4 | 5.765 | 5.765 | 3.1
5 | 7.2 | 7.199 | 2.2
rick | 2024-12-19 09:38
Index | Set Point | Laser | Measure |  | Slope | Offset | Constant
1 | 1.1 | 1.102 | 8.485 |  | -0.994 | 9.625 | -0.994
2 | 2.672 | 2.671 | 7.07
3 | 4.214 | 4.213 | 5.397
4 | 5.688 | 5.687 | 3.93
5 | 7.2 | 7.2 | 2.491
rick A | 2024-12-19 13:51
Index | Set Point | Laser | Measure |  | Slope | Offset | Constant
1 | 1.1 | 1.102 | 8.499 |  | -1.007 | 9.628 | -1.007
2 | 2.659 | 2.659 | 6.956
3 | 4.218 | 4.218 | 5.396
4 | 5.702 | 5.701 | 3.923
5 | 7.2 | 7.199 | 2.345
HLC | 2024-12-19 14:12
Index | Set Point | Laser | Measure |  | Slope | Offset | Constant
1 | 1.1 | 1.102 | 8.508 |  | -0.997 | 9.61 | -0.997
2 | 2.628 | 2.625 | 6.995
3 | 4.199 | 4.198 | 5.424
4 | 5.719 | 5.719 | 3.908
5 | 7.2 | 7.199 | 2.427
HLC | 2024-12-19 14:22
Index | Set Point | Laser | Measure |  | Slope | Offset | Constant
1 | 1.1 | 1.102 | 8.508 |  | -0.996 | 9.608 | -0.996
2 | 2.633 | 2.631 | 6.987
3 | 4.154 | 4.154 | 5.471
4 | 5.678 | 5.677 | 3.957
5 | 7.2 | 7.2 | 2.429
HLC | 2024-12-19 14:31
Index | Set Point | Laser | Measure |  | Slope | Offset | Constant
1 | 1.1 | 1.102 | 8.506 |  | -0.996 | 9.607 | -0.996
2 | 2.638 | 2.637 | 6.984
3 | 4.214 | 4.212 | 5.413
4 | 5.683 | 5.682 | 3.942
5 | 7.2 | 7.199 | 2.434
DIANE | 2024-12-19 14:40
Index | Set Point | Laser | Measure |  | Slope | Offset | Constant
1 | 1.1 | 1.102 | 8.559 |  | -1.001 | 9.657 | -1.001
2 | 2.725 | 2.724 | 6.916
3 | 4.171 | 4.171 | 5.477
4 | 5.701 | 5.7 | 3.974
5 | 7.2 | 7.199 | 2.437
rwh | 2024-12-19 15:00
Index | Set Point | Laser | Measure |  | Slope | Offset | Constant
1 | 1.1 | 1.102 | 8.493 |  | -0.996 | 9.592 | -0.996
2 | 2.638 | 2.636 | 6.968
3 | 4.247 | 4.244 | 5.365
4 | 5.688 | 5.686 | 3.927
5 | 7.2 | 7.2 | 2.421
RWH | 2024-12-19 15:11
Index | Set Point | Laser | Measure |  | Slope | Offset | Constant
1 | 1.1 | 1.102 | 8.492 |  | -0.994 | 9.585 | -0.994
2 | 2.7 | 2.698 | 6.903
3 | 4.158 | 4.158 | 5.452
4 | 5.717 | 5.717 | 3.9
5 | 7.2 | 7.2 | 2.434
RWH | 2024-12-19 15:32
Index | Set Point | Laser | Measure |  | Slope | Offset | Constant
1 | 1.1 | 1.102 | 8.493 |  | -0.995 | 9.592 | -0.995
2 | 2.694 | 2.694 | 6.91
3 | 4.226 | 4.234 | 5.385
4 | 5.747 | 5.746 | 3.875
5 | 7.2 | 7.198 | 2.425
hlc | 2024-12-20 08:07
Index | Set Point | Laser | Measure |  | Slope | Offset | Constant
1 | 1.1 | 1.102 | 8.509 |  | -0.995 | 9.609 | -0.995
2 | 2.71 | 2.71 | 6.917
3 | 4.222 | 4.221 | 5.405
4 | 5.733 | 5.732 | 3.902
5 | 7.2 | 7.199 | 2.446
hlc | 2024-12-20 08:16
Index | Set Point | Laser | Measure |  | Slope | Offset | Constant
1 | 1.1 | 1.101 | 8.514 |  | -0.997 | 9.612 | -0.997
2 | 2.643 | 2.642 | 6.98
3 | 4.218 | 4.217 | 5.409
4 | 5.761 | 5.761 | 3.869
5 | 7.2 | 7.199 | 2.437
diane | 2024-12-20 08:24
Index | Set Point | Laser | Measure |  | Slope | Offset | Constant
1 | 1.1 | 1.102 | 8.581 |  | -0.998 | 9.374 | -0.998
2 | 2.699 | 2.697 | 6.922
3 | 4.162 | 4.162 | 4.115
4 | 5.715 | 5.715 | 3.95
5 | 7.2 | 7.199 | 2.471
diane | 2024-12-20 08:38
Index | Set Point | Laser | Measure |  | Slope | Offset | Constant
1 | 1.1 | 1.101 | 8.52 |  | -0.988 | 9.583 | -0.988
2 | 2.693 | 2.694 | 6.911
3 | 4.189 | 4.189 | 5.401
4 | 5.711 | 5.711 | 3.97
5 | 7.2 | 7.199 | 2.478
diane | 2024-12-20 08:49
Index | Set Point | Laser | Measure |  | Slope | Offset | Constant
1 | 1.1 | 1.102 | 8.523 |  | -0.852 | 8.305 | -0.852
2 | 2.706 | 2.704 | 4.75
3 | 4.195 | 4.195 | 3.987
4 | 5.726 | 5.725 | 3.98
5 | 7.2 | 7.2 | 2.462
diane | 2024-12-20 08:59
Index | Set Point | Laser | Measure |  | Slope | Offset | Constant
1 | 1.1 | 1.102 | 8.516 |  | -0.991 | 9.6 | -0.991
2 | 2.716 | 2.715 | 6.916
3 | 4.242 | 4.241 | 5.371
4 | 5.717 | 5.716 | 3.934
5 | 7.2 | 7.2 | 2.477
diane | 2024-12-20 09:09
Index | Set Point | Laser | Measure |  | Slope | Offset | Constant
1 | 1.1 | 1.101 | 8.525 |  | -0.999 | 10.134 | -0.999
2 | 2.652 | 2.651 | 6.958
3 | 4.157 | 4.156 | 8.024
4 | 5.774 | 5.774 | 3.853
5 | 7.2 | 7.199 | 2.44
RWH | 2024-12-20 09:26
Index | Set Point | Laser=X | Measure |  | Slope=M | Offset=B | Constant
1 | 1.1 | 1.102 | 8.495 |  | -0.996 | 9.594 | -0.996
2 | 2.714 | 2.712 | 6.896
3 | 4.197 | 4.196 | 5.417 | M*x + B | 8.4964 | -0.001408
4 | 5.697 | 5.696 | 3.922
5 | 7.2 | 7.199 | 2.423
RWH | 2024-12-20 09:38
Index | Set Point | Laser = X | Measure =Y |  | Slope | Offset | Constant
1 | 1.1 | 1.101 | 8.497 | 8.5004 | -0.996 | 9.597 | -0.996
2 | 2.67 | 2.67 | 6.941 | 6.9377
3 | 4.158 | 4.157 | 5.456 | 5.4566
4 | 5.691 | 5.69 | 3.927 | 3.9298
5 | 7.2 | 7.199 | 2.425 | 2.4268
|  | 4.1634 | 5.4492
9.6126
Note on H228: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: There must have been a bug in as the slope is recorded in the Constant column.

## Processed
Trial | Date | Operator | Slope | Offset |  |  | Constant (slope = -1)
1 | 2024-11-06 13:17 | Dragan | -0.996 | 9.591
2 | 2024-11-06 13:27 | George | -0.996 | 9.595
3 | 2024-11-06 13:40 | Dragan | -0.995 | 9.601
4 | 2024-11-06 13:49 | George | -0.996 | 9.601
5 | 2024-11-06 13:59 | Dragan | -0.996 | 9.602
6 | 2024-12-19 09:38 | rick A | -0.994 | 9.625 |  |  | -0.994
7 | 2024-12-19 13:51 | rick A | -1.007 | 9.628 |  |  | -1.007
19 | 2024-12-20 08:49 | DIANE | -0.852 | 8.305 | Outlier |  | -0.852
|  | Average | -0.99567 | 9.6236 | Slope Deviation | Offset Deviation
8 | 2024-12-19 14:12 | HLC | -0.997 | 9.61 | 0.0013333 | 0.0136 | -0.997
9 | 2024-12-19 14:22 | HLC | -0.996 | 9.608 | 0.00033333 | 0.0156 | -0.996
10 | 2024-12-19 14:31 | HLC | -0.996 | 9.607 | 0.00033333 | 0.0166 | -0.996
11 | 2024-12-19 14:40 | DIANE | -1.001 | 9.657 | 0.0053333 | -0.0334 | -1.001
12 | 2024-12-19 15:00 | RWH | -0.996 | 9.592 | 0.00033333 | 0.0316 | -0.996
13 | 2024-12-19 15:11 | RWH | -0.994 | 9.585 | -0.0016667 | 0.0386 | -0.994
14 | 2024-12-19 15:32 | RWH | -0.995 | 9.592 | -0.00066667 | 0.0316 | -0.995
15 | 2024-12-20 08:07 | HLC | -0.995 | 9.609 | -0.00066667 | 0.0146 | -0.995
16 | 2024-12-20 08:16 | HLC | -0.997 | 9.612 | 0.0013333 | 0.0116 | -0.997
17 | 2024-12-20 08:24 | DIANE | -0.998 | 9.374 | 0.0023333 | 0.2496 | -0.998
18 | 2024-12-20 08:38 | DIANE | -0.988 | 9.583 | -0.0076667 | 0.0406 | -0.988
20 | 2024-12-20 08:59 | DIANE | -0.991 | 9.6 | -0.0046667 | 0.0236 | -0.991
21 | 2024-12-20 09:09 | DIANE | -0.999 | 10.134 | 0.0033333 | -0.5104 | -0.999
22 | 2024-12-20 09:26 | RWH | -0.996 | 9.594 | 0.00033333 | 0.0296 | -0.996
23 | 2024-12-20 09:38 | RWH | -0.996 | 9.597 | 0.00033333 | 0.0266 | -0.996
This chart isn't available in your version of Excel. Editing this shape or saving this workbook into a different file format will permanently break the chart.
This chart isn't available in your version of Excel. Editing this shape or saving this workbook into a different file format will permanently break the chart.
This chart isn't available in your version of Excel. Editing this shape or saving this workbook into a different file format will permanently break the chart.
This chart isn't available in your version of Excel. Editing this shape or saving this workbook into a different file format will permanently break the chart.

## Processed skilled only
Trial | Date | Operator | Slope | Offset
1 | 2024-11-06 13:17 | Dragan | -0.996 | 9.591
2 | 2024-11-06 13:27 | George | -0.996 | 9.595
3 | 2024-11-06 13:40 | Dragan | -0.995 | 9.601
4 | 2024-11-06 13:49 | George | -0.996 | 9.601
5 | 2024-11-06 13:59 | Dragan | -0.996 | 9.602
6 | 2024-12-19 09:38 | rick A | -0.994 | 9.625
7 | 2024-12-19 13:51 | rick A | -1.007 | 9.628
19 | 2024-12-20 08:49 | DIANE | -0.852 | 8.305 | Outlier
11 | 2024-12-19 14:40 | DIANE | -1.001 | 9.657 | 0.0052 | -0.0564
17 | 2024-12-20 08:24 | DIANE | -0.998 | 9.374 | 0.0022 | 0.2266
18 | 2024-12-20 08:38 | DIANE | -0.988 | 9.583 | -0.0078 | 0.0176
20 | 2024-12-20 08:59 | DIANE | -0.991 | 9.6 | -0.0048 | 0.0006
21 | 2024-12-20 09:09 | DIANE | -0.999 | 10.134 | 0.0032 | -0.5334
Trial | Date | Operator | Slope | Offset | Slope Deviation | Offset Deviation
8 | 2024-12-19 14:12 | HLC | -0.997 | 9.61 | 0.0012 | -0.0094
9 | 2024-12-19 14:22 | HLC | -0.996 | 9.608 | 0.0002 | -0.0074
10 | 2024-12-19 14:31 | HLC | -0.996 | 9.607 | 0.0002 | -0.0064
15 | 2024-12-20 08:07 | HLC | -0.995 | 9.609 | -0.0008 | -0.0084
16 | 2024-12-20 08:16 | HLC | -0.997 | 9.612 | 0.0012 | -0.0114
12 | 2024-12-19 15:00 | RWH | -0.996 | 9.592 | 0.0002 | 0.0086
13 | 2024-12-19 15:11 | RWH | -0.994 | 9.585 | -0.0018 | 0.0156
14 | 2024-12-19 15:32 | RWH | -0.995 | 9.592 | -0.0008 | 0.0086
22 | 2024-12-20 09:26 | RWH | -0.996 | 9.594 | 0.0002 | 0.0066
23 | 2024-12-20 09:38 | RWH | -0.996 | 9.597 | 0.0002 | 0.0036
|  | Max |  |  | 0.0012 | 0.0156
|  | Min |  |  | -0.0018 | -0.0114
|  | Range |  |  | 0.003 | 0.027
|  |  | Slope | Offset | Slope CI upper | Slope CI lower | Offset CI upper | Offset CI lower |  | Expected | Upper | Lower
|  | Average | -0.9958 | 9.6006 | -0.9957 | -0.9959 | 9.6016 | 9.5996 | Slope | -0.9958 | -0.9946 | -0.99695 | F38 | G36
|  | StdDev | 0.00087178 | 0.0091236 |  |  |  |  | Offset | 9.6006 | 9.6109 | 9.5881 | H36 | I38
[9 empty] | Index |  |  |  | Upper | Lower
|  | Expected Zero | Slope | Offset | Slope UL /Dev | Slope LL /Dev | Offset UL /Dev | Offset LL /Dev | 2 | 7.609 | 7.6217 | 7.5942 | -0.01272 | 0.014846
|  | Average HLC | -0.9962 | 9.6092 | -0.99545 | -0.99695 | 9.6109 | 9.6075 | 4 | 5.6174 | 5.6325 | 5.6003 | -0.01512 | 0.017143
|  | StdDev HLC | 0.00074833 | 0.0017205 | -0.00024973 | 0.0010497 | -0.0092886 | -0.0079114 | 6 | 3.6258 | 3.6433 | 3.6064 | -0.01752 | 0.01944
|  | Average RWH | -0.9954 | 9.592 | -0.9946 | -0.9962 | 9.5959 | 9.5881 | 8 | 1.6342 | 1.6541 | 1.6125 | -0.01992 | 0.021736
|  | StdDev RWH | 0.0008 | 0.0039497 | -0.0011014 | 0.0003014 | 0.0056822 | 0.011518
|  | α = 0.05 2-Tail | T-Stat | Num of StDev= | 1
|  | df = 4 | 2.776
|  | df = 9 | 2.262
Operators
RWH Rob Hyde
DIANE Dian Daugherty
HLC Hannah Chafin
Each operator made 5 measurements trials of 5-points on the LV automated test rig using OLS_LevelZeroVerification.vi 12/18/24. The first and fifth points had fixed set point value, but observations 3-4 had a random amount +/-0.100" added to the test set point value to achieve a "Blind" observation at those points.
The evaluation was concuded over two days, 12/19/24 and 12/20/24.
A 5-point regression is calculated using MS Excel Analysis Tool Pack on the 5 raw observations for each trial. The regression slope and Offset are considered the results of a trialSlope and Offset = Results of 5 point regression Analysis
Deviation = Average - Observation
Tirals 1-8 are previous exploritory evaluations.
I removed Operator DIANE as "Unskilled" due to size of varaiation in measumrents.
The Measuremnet system was differnece between Set Point and Laser measurment was biased hight 0.002" with StDev of 0.0014 over 107 observations.
C I = Xavg+/- T-Stat * StDev / √N for @95% certainty α = 0.05 2-Tailed
ref: 0.005" is 0.0625% of 8"
This chart isn't available in your version of Excel. Editing this shape or saving this workbook into a different file format will permanently break the chart.
This chart isn't available in your version of Excel. Editing this shape or saving this workbook into a different file format will permanently break the chart.
This chart isn't available in your version of Excel. Editing this shape or saving this workbook into a different file format will permanently break the chart.
This chart isn't available in your version of Excel. Editing this shape or saving this workbook into a different file format will permanently break the chart.