## Elevate Level
Dragan | 2024-11-06 13:17
Index | Set Point | Laser | Measure | Slope | Offset
1 | 1.1 | 1.102 | 8.495 | -0.996 | 9.591
2 | 2.7 | 2.664 | 6.938
3 | 4.3 | 4.26 | 5.347
4 | 5.8 | 5.756 | 3.857
5 | 7.2 | 7.249 | 2.374
George | 2024-11-06 13:27
Index | Set Point | Laser | Measure | Slope | Offset
1 | 1.1 | 1.102 | 8.499 | -0.996 | 9.595
2 | 2.7 | 2.735 | 6.876
3 | 4.2 | 4.188 | 5.417
4 | 5.8 | 5.793 | 3.822
5 | 7.2 | 7.248 | 2.384
Dragan | 2024-11-06 13:40
Index | Set Point | Laser | Measure | Slope | Offset
1 | 1.1 | 1.102 | 8.5 | -0.995 | 9.601
2 | 2.7 | 2.722 | 6.897
3 | 4.2 | 4.213 | 5.406
4 | 5.7 | 5.738 | 3.888
5 | 7.2 | 7.249 | 2.384
George | 2024-11-06 13:49
Index | Set Point | Laser | Measure | Slope | Offset
1 | 1.1 | 1.102 | 8.503 | -0.996 | 9.601
2 | 2.7 | 2.703 | 6.909
3 | 4.2 | 4.176 | 5.444
4 | 5.7 | 5.745 | 3.881
5 | 7.2 | 7.248 | 2.382
Dragan | 2024-11-06 13:59
Index | Set Point | Laser | Measure | Slope | Offset
1 | 1.1 | 1.102 | 8.504 | -0.996 | 9.602
2 | 2.7 | 2.676 | 6.936
3 | 4.2 | 4.178 | 5.442
4 | 5.7 | 5.715 | 3.908
5 | 7.2 | 7.249 | 2.381
8.504 | 2024-11-06 14:09
Index | Set Point | Laser | Measure | Slope | Offset
1 | 1.1 | 1.101 | 8.504 | -0.996 | 9.602
2 | 2.7 | 2.707 | 6.907
3 | 4.2 | 4.183 | 5.439
4 | 5.7 | 5.716 | 3.906
5 | 7.2 | 7.249 | 2.383
rick | 2024-11-17 09:07
Index | Set Point | Laser | Measure | Slope | Offset
1 | 1 | 1.002 | 0 | 0 | 0
2 | 2.5 | 2.53 | 0
3 | 4.1 | 4.092 | 0
4 | 5.5 | 5.502 | 0
5 | 7 | 7.002 | 0
george | 45613.6514
Index | Set Point | Laser | Measure | Slope | Offset
1 | 1.1 | 1.102 | 0 | 0 | 0
2 | 2.7 | 2.66 | 0
3 | 4.2 | 4.248 | 0
4 | 5.7 | 5.7 | 0
5 | 7.2 | 7.202 | 0
rob | 2024-11-17 15:56
Index | Set Point | Laser | Measure | Slope | Offset
1 | 1.1 | 1.102 | 0 | 0 | 0
2 | 2.723 | 2.724 | 0
3 | 4.175 | 4.174 | 0
4 | 5.683 | 5.682 | 0
5 | 7.2 | 7.204 | 0
2024-11-17 16:29
Index | Set Point | Laser | Measure | Slope | Offset
1 | 1.1 | 1.102 | 8 | -1.003 | 9.066
2 | 2.67 | 2.67 | 6.5
3 | 4.165 | 4.166 | 4.7
4 | 5.725 | 5.728 | 3.2
5 | 7.2 | 7.202 | 2
rick | 2025-05-07 09:26
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.206 | 8.403 | 0.002 | -0.992 | 9.598
2 | 7.032 | 2.575 | 7.044
3 | 5.65 | 3.963 | 5.666
4 | 4.231 | 5.393 | 4.249
5 | 2.75 | 6.88 | 2.775
George | 2025-05-07 13:17
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.206 | 8.388 | 0.012 | -0.996 | 9.595
2 | 7.055 | 2.554 | 7.05
3 | 5.632 | 3.987 | 5.635
4 | 4.252 | 5.371 | 4.245
5 | 2.75 | 6.875 | 2.739
Camren | 2025-05-07 13:28
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.207 | 8.397 | 0.763 | -0.859 | 8.805
2 | 6.998 | 1.21 | 7.003
3 | 5.615 | 4.002 | 5.614
4 | 4.164 | 5.458 | 4.16
5 | 2.75 | 6.879 | 2.747
Camren | 2025-05-07 14:02
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.207 | 8.377 | 1.135 | -1.012 | 9.359
2 | 7.044 | 2.568 | 7.021
3 | 5.671 | 3.944 | 5.657
4 | 4.255 | 3.943 | 4.232
5 | 2.75 | 6.88 | 2.74
Camren | 2025-05-07 14:19
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.207 | 8.393 | 0.007 | -0.996 | 9.594
2 | 7.031 | 2.575 | 7.023
3 | 5.63 | 3.989 | 5.627
4 | 4.205 | 5.413 | 4.2
5 | 2.75 | 6.878 | 2.74
Camren | 2025-05-07 14:31
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.207 | 8.392 | 0.008 | -0.996 | 9.595
2 | 7.057 | 2.555 | 7.058
3 | 5.585 | 4.028 | 5.577
4 | 4.25 | 5.371 | 4.253
5 | 2.75 | 6.877 | 2.748
George | 2025-05-07 15:29
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.207 | 8.383 | 0.763 | -1.354 | 9.863
2 | 7.035 | 2.58 | 7.035
3 | 5.629 | 2.571 | 5.617
4 | 4.21 | 3.987 | 4.204
5 | 2.75 | 5.41 | 2.738
George | 2025-05-07 15:46
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.209 | 8.379 | 0.004 | -0.994 | 9.585
2 | 7.067 | 2.544 | 7.059
3 | 5.622 | 3.996 | 5.615
4 | 4.239 | 5.383 | 4.234
5 | 2.75 | 6.879 | 2.743
George | 2025-05-07 16:03
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.205 | 8.388 | 0.005 | -0.996 | 9.591
2 | 6.991 | 2.619 | 6.982
3 | 5.65 | 3.964 | 5.641
4 | 4.209 | 5.41 | 4.205
5 | 2.75 | 6.876 | 2.736
George | 2025-05-08 07:42
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.208 | 8.384 | 0.004 | -0.997 | 9.593
2 | 7.06 | 2.552 | 7.051
3 | 5.668 | 3.947 | 5.658
4 | 4.251 | 5.369 | 4.24
5 | 2.75 | 6.871 | 2.736
George | 2025-05-08 07:57
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.206 | 8.388 | 0.003 | -0.996 | 9.591
2 | 7.062 | 2.55 | 7.053
3 | 5.595 | 4.019 | 5.587
4 | 4.18 | 5.443 | 4.17
5 | 2.75 | 6.88 | 2.736
Rick for Reference | 2025-05-08 10:57
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.209 | 8.403 | 0.004 | -0.992 | 9.601
2 | 7.08 | 2.528 | 7.088
3 | 5.611 | 4.007 | 5.627
4 | 4.237 | 5.386 | 4.254
5 | 2.75 | 6.879 | 2.775
Rick for Reference | 2025-05-08 11:10
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.206 | 8.405 | 0.003 | -0.992 | 9.601
2 | 7.012 | 2.597 | 7.023
3 | 5.66 | 3.956 | 5.676
4 | 4.207 | 5.416 | 4.224
5 | 2.75 | 6.88 | 2.775
Gabby | 2025-05-08 13:57
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.209 | 8.384 | 0.987 | -0.925 | 9.064
2 | 7 | 2.612 | 7
3 | 5.665 | 2.61 | 5.662
4 | 4.241 | 5.382 | 4.234
5 | 2.75 | 6.879 | 2.747
Gabby | 2025-05-08 14:10
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.206 | 8.387 | 0.008 | -0.995 | 9.593
2 | 7.068 | 2.544 | 7.067
3 | 5.589 | 4.029 | 5.581
4 | 4.187 | 5.437 | 4.183
5 | 2.75 | 6.88 | 2.741
Gabby | 2025-05-08 14:21
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.207 | 8.384 | 0.965 | -1.121 | 9.136
2 | 6.997 | 1.213 | 6.99
3 | 5.612 | 4.004 | 5.612
4 | 4.25 | 4.003 | 4.243
5 | 2.75 | 5.371 | 2.74
Dragan | 2025-05-08 14:34
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.207 | 8.386 | 1.054 | -0.568 | 7.125
2 | 7.065 | 1.211 | 7.063
3 | 5.653 | 3.965 | 5.643
4 | 4.163 | 5.462 | 4.155
5 | 2.75 | 6.879 | 2.746
Dragan | 2025-05-08 14:42
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.207 | 8.387 | 0.005 | -0.994 | 9.589
2 | 6.99 | 2.621 | 6.986
3 | 5.631 | 3.986 | 5.623
4 | 4.223 | 5.401 | 4.225
5 | 2.75 | 6.881 | 2.745
Dragan | 2025-05-08 14:50
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.207 | 8.386 | 0.01 | -0.995 | 9.596
2 | 7.078 | 2.53 | 7.087
3 | 5.669 | 3.949 | 5.667
4 | 4.205 | 5.419 | 4.205
5 | 2.75 | 6.879 | 2.743
rick for reference | 2025-05-13 16:12
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.208 | 8.404 | 0 | -0.993 | 9.603
2 | 7.001 | 2.608 | 7.013
3 | 5.671 | 3.945 | 5.685
4 | 4.255 | 5.368 | 4.272
5 | 2.75 | 6.879 | 2.773
rick for reference | 2025-05-13 16:22
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.202 | 8.408 | 0 | -0.993 | 9.602
2 | 7.033 | 2.579 | 7.042
3 | 5.671 | 3.944 | 5.686
4 | 4.223 | 5.398 | 4.242
5 | 2.75 | 6.878 | 2.773
George | 2025-05-14 09:31
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.208 | 8.393 | 0.008 | -0.997 | 9.599
2 | 7.002 | 2.608 | 6.997
3 | 5.597 | 4.018 | 5.602
4 | 4.262 | 5.358 | 4.256
5 | 2.75 | 6.878 | 2.742
George | 2025-05-14 09:43
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.207 | 8.395 | 0.005 | -0.997 | 9.598
2 | 6.99 | 2.62 | 6.987
3 | 5.613 | 3.999 | 5.611
4 | 4.205 | 5.413 | 4.208
5 | 2.75 | 6.875 | 2.743
George | 2025-05-14 13:42
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.207 | 8.399 | 0.008 | -0.998 | 9.606
2 | 7.083 | 2.53 | 7.08
3 | 5.666 | 3.95 | 5.673
4 | 4.164 | 5.458 | 4.161
5 | 2.75 | 6.874 | 2.743
Camren | 2025-05-14 14:03
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.21 | 8.39 | 0.003 | -0.997 | 9.599
2 | 7.06 | 2.55 | 7.059
3 | 5.638 | 3.972 | 5.64
4 | 4.243 | 5.37 | 4.24
5 | 2.75 | 6.878 | 2.74
Camren | 2025-05-14 14:18
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.211 | 8.383 | 0.013 | -0.997 | 9.594
2 | 7.064 | 2.547 | 7.052
3 | 5.63 | 3.984 | 5.636
4 | 4.176 | 5.444 | 4.167
5 | 2.75 | 6.875 | 2.738
Camren | 2025-05-14 14:28
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.208 | 8.392 | 0.008 | -0.997 | 9.594
2 | 7.033 | 2.57 | 7.034
3 | 5.628 | 3.986 | 5.614
4 | 4.256 | 5.361 | 4.251
5 | 2.75 | 6.876 | 2.744
rick for reference | 2025-05-15 14:52
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.206 | 8.406 | 0 | -0.993 | 9.603
2 | 7.074 | 2.524 | 7.096
3 | 5.656 | 3.959 | 5.671
4 | 4.186 | 5.437 | 4.204
5 | 2.75 | 6.877 | 2.774
Gabby | 2025-05-19 09:47
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.209 | 8.388 | 0.003 | -0.997 | 9.596
2 | 7.052 | 2.55 | 7.056
3 | 5.622 | 3.993 | 5.618
4 | 4.253 | 5.368 | 4.242
5 | 2.75 | 6.875 | 2.74
Gabby | 2025-05-19 09:58
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.21 | 8.39 | 0.015 | -0.993 | 9.587
2 | 7.066 | 2.54 | 7.049
3 | 5.618 | 3.999 | 5.62
4 | 4.245 | 5.376 | 4.26
5 | 2.75 | 6.879 | 2.744
Gabby | 2025-05-19 10:08
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.209 | 8.393 | 0.005 | -0.997 | 9.6
2 | 7.038 | 2.571 | 7.042
3 | 5.66 | 3.956 | 5.654
4 | 4.167 | 5.456 | 4.16
5 | 2.75 | 6.878 | 2.743
Dragan | 2025-05-19 10:16
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.208 | 8.39 | 0.002 | -0.995 | 9.59
2 | 6.998 | 2.613 | 6.99
3 | 5.618 | 3.993 | 5.617
4 | 4.167 | 5.454 | 4.165
5 | 2.75 | 6.883 | 2.745
Dragan | 2025-05-19 10:26
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.209 | 8.39 | 0.006 | -0.996 | 9.595
2 | 7.067 | 2.543 | 7.06
3 | 5.596 | 4.015 | 5.596
4 | 4.227 | 5.395 | 4.228
5 | 2.75 | 6.871 | 2.746
Dragan | 2025-05-19 10:51
Index | Set Point | Laser | Measure | Goodness | Slope | Intercept
1 | 8.4 | 1.208 | 8.392 | 0.005 | -0.998 | 9.6
2 | 7.009 | 2.599 | 7.012
3 | 5.637 | 3.974 | 5.636
4 | 4.214 | 5.407 | 4.205
5 | 2.75 | 6.88 | 2.736

## Elevate Resistance
Rick A | 2025-04-09 12:52 | Tol = +/- 1LSD
Index | Ref (W) Default | Fluke Meter (W) | DataQ (Volts) |  | Slope | Offset | Constant
R100-A | 100.07 | 100.07 | 0.8893 | ChA | ChA | 112.456 | 0.062918
R100-B | 100.08 | 100.08 | 0.8802 | ChB | ChB | 113.8686 | -0.14715
R1K-A | 1000.1 | 1000.1 | 8.8927
R1K-B | 1000.2 | 1000.2 | 8.7851
rick | 2025-05-07 09:25 | Default
Index | Reference | Fluke Meter | DataQ |  | Channel | Slope | Intercept
R100-A | 100.08 | 100.08 | 0.889 |  | Ch A | 112.484 | 0.061
R100-B | 100.08 | 100.08 | 0.88 |  | Ch B | 113.938 | -1.232
R1K-A | 1000.2 | 1000.2 | 8.891
R1K-B | 1000.2 | 1000.2 | 8.789
Camren | 2025-05-07 13:46 | Default
Index | Reference | Fluke Meter | DataQ |  | Channel | Slope | Intercept
R100-A | 100.08 | 100.08 | 0.889 |  | Ch A | 112.459 | 0.109
R100-B | 100.08 | 100.08 | 0.88 |  | Ch B | 113.933 | -1.202
R1K-A | 1000.2 | 1000.2 | 8.893
R1K-B | 1000.2 | 1000.2 | 8.789
rick | 2025-05-08 10:07 | Default
Index | Reference | Fluke Meter | DataQ |  | Channel | Slope | Intercept
R100-A | 100.08 | 100.08 | 0.889 |  | Ch A | 112.462 | 0.074
R100-B | 100.08 | 100.08 | 0.88 |  | Ch B | 113.979 | -1.275
R1K-A | 1000.2 | 1000.2 | 8.893
R1K-B | 1000.2 | 1000.2 | 8.786

## Level R&R
CAL ID | Operator | Date/Time | OP | Trial | Goodness | Slope | Intercept | Slope Dev | Int Dev | [5 empty] | Count |  |  | CL | Z-score
1 | George | 2025-05-14 09:31 | 1 | 1 | 0.008 | -0.997 | 9.599 | 0.00041667 | 0.0025 |  |  |  |  | N | 12 |  |  | 90% | 1.645 | 2.58
2 | George | 2025-05-14 09:43 | 1 | 2 | 0.005 | -0.997 | 9.598 | 0.00041667 | 0.0015 |  |  |  |  | NTesters | 4 |  |  | 96% | 1.96
3 | George | 2025-05-14 13:42 | 1 | 3 | 0.008 | -0.998 | 9.606 | 0.0014167 | 0.0095 |  |  |  |  | NReplicates | 2 |  |  | 99% | 2.58
4 | Camren | 2025-05-14 14:03 | 2 | 1 | 0.003 | -0.997 | 9.599 | 0.00041667 | 0.0025
5 | Camren | 2025-05-14 14:18 | 2 | 2 | 0.013 | -0.997 | 9.594 | 0.00041667 | 0.0025 |  | exclude-> |  |  |  | Slope | Intercept |  | Level | LCLmin | LCLmax | MUI* | ± inches
6 | Camren | 2025-05-14 14:28 | 3 | 3 | 0.008 | -0.997 | 9.594 | 0.00041667 | 0.0025 |  |  |  |  | Average | -0.99658 | 9.5965 |  | 8.4 | -8.3799 | -8.3627 | -0.017254 | 0.008627
7 | Gabby | 2025-05-19 09:47 | 3 | 1 | 0.003 | -0.997 | 9.596 | 0.00041667 | 0.0005 |  | CL |  |  | Std Dev | 0.001379 | 0.0050181 |  | 2.75 | -2.7434 | -2.7378 | -0.0056486 | 0.0028243
8 | Gabby | 2025-05-19 09:58 | 3 | 2 | 0.015 | -0.993 | 9.587 | 0.0035833 | 0.0095 |  | 99% |  |  | UCL | -0.99761 | 9.5928
9 | Gabby | 2025-05-19 10:08 | 3 | 3 | 0.005 | -0.997 | 9.6 | 0.00041667 | 0.0035 |  |  |  |  | LCl | -0.99556 | 9.6002 |  |  |  |  | ±Tolerance = | ±0.103%
10 | Dragan | 2025-05-19 10:16 | 4 | 1 | 0.002 | -0.995 | 9.59 | 0.0015833 | 0.0065
11 | Dragan | 2025-05-19 10:16 | 4 | 2 | 0.006 | -0.996 | 9.595 | 0.00058333 | 0.0015
12 | Dragan | 2025-05-19 10:51 | 4 | 3 | 0.005 | -0.998 | 9.6 | 0.0014167 | 0.0035
Run ID | Operator | Date/Time | Op | Trial | Set Point | Set Point2 | Op * Laser | Laser^2 | OP(2) | Laser | Measure | Column14 | Column1 | Column2 | Column3 | Column4 | Column5 | Column6 | Column7 | Column8 | Column9 | Column10 | Column11 | Column12 | Column13
1 | George | 2025-05-14 09:31 | 1 | 1 | 1 | 8.4 | 1.4593 | 1.4593 | 1 | 1.208 | 8.393
2 | George | 2025-05-14 09:31 | 1 | 1 | 2 | 7.002 | 6.8017 | 6.8017 | 1 | 2.608 | 6.997
3 | George | 2025-05-14 09:31 | 1 | 1 | 3 | 5.597 | 16.1443 | 16.1443 | 1 | 4.018 | 5.602
4 | George | 2025-05-14 09:31 | 1 | 1 | 4 | 4.262 | 28.7082 | 28.7082 | 1 | 5.358 | 4.256
5 | George | 2025-05-14 09:31 | 1 | 1 | 5 | 2.75 | 47.3069 | 47.3069 | 1 | 6.878 | 2.742
6 | George | 2025-05-14 09:43 | 1 | 2 | 1 | 8.4 | 1.4568 | 1.4568 | 1 | 1.207 | 8.395
7 | George | 2025-05-14 09:43 | 1 | 2 | 2 | 6.99 | 6.8644 | 6.8644 | 1 | 2.62 | 6.987
8 | George | 2025-05-14 09:43 | 1 | 2 | 3 | 5.613 | 15.992 | 15.992 | 1 | 3.999 | 5.611
9 | George | 2025-05-14 09:43 | 1 | 2 | 4 | 4.205 | 29.3006 | 29.3006 | 1 | 5.413 | 4.208 |  | SUMMARY OUTPUT
10 | George | 2025-05-14 09:43 | 1 | 2 | 5 | 2.75 | 47.2656 | 47.2656 | 1 | 6.875 | 2.743
11 | George | 2025-05-14 13:42 | 1 | 3 | 1 | 8.4 | 1.4568 | 1.4568 | 1 | 1.207 | 8.399 |  | Regression Statistics
12 | George | 2025-05-14 13:42 | 1 | 3 | 2 | 7.083 | 6.4009 | 6.4009 | 1 | 2.53 | 7.08 |  | Multiple R | 1
13 | George | 2025-05-14 13:42 | 1 | 3 | 3 | 5.666 | 15.6025 | 15.6025 | 1 | 3.95 | 5.673 |  | R Square | 0.99999
14 | George | 2025-05-14 13:42 | 1 | 3 | 4 | 4.164 | 29.7898 | 29.7898 | 1 | 5.458 | 4.161 |  | Adjusted R Square | 0.99999
15 | George | 2025-05-14 13:42 | 1 | 3 | 5 | 2.75 | 47.2519 | 47.2519 | 1 | 6.874 | 2.743 |  | Standard Error | 0.005648
16 | Camren | 2025-05-14 14:03 | 2 | 1 | 1 | 8.4 | 2.9282 | 1.4641 | 2 | 1.21 | 8.39 |  | Observations | 60
17 | Camren | 2025-05-14 14:03 | 2 | 1 | 2 | 7.06 | 13.005 | 6.5025 | 2 | 2.55 | 7.059
18 | Camren | 2025-05-14 14:03 | 2 | 1 | 3 | 5.638 | 31.5536 | 15.7768 | 2 | 3.972 | 5.64 |  | ANOVA
19 | Camren | 2025-05-14 14:03 | 2 | 1 | 4 | 4.243 | 57.6738 | 28.8369 | 2 | 5.37 | 4.24 |  |  | df | SS | MS | F | Significance F
20 | Camren | 2025-05-14 14:03 | 2 | 1 | 5 | 2.75 | 94.6138 | 47.3069 | 2 | 6.878 | 2.74 |  | Regression | 4 | 239.3772 | 59.8443 | 1876008.6134 | 1.7544e-140
21 | Camren | 2025-05-14 14:18 | 2 | 2 | 1 | 8.4 | 2.933 | 1.4665 | 2 | 1.211 | 8.383 |  | Residual | 55 | 0.0017545 | 3.19e-05
22 | Camren | 2025-05-14 14:18 | 2 | 2 | 2 | 7.064 | 12.9744 | 6.4872 | 2 | 2.547 | 7.052 |  | Total | 59 | 239.3789
23 | Camren | 2025-05-14 14:18 | 2 | 2 | 3 | 5.63 | 31.7445 | 15.8723 | 2 | 3.984 | 5.636
24 | Camren | 2025-05-14 14:18 | 2 | 2 | 4 | 4.176 | 59.2743 | 29.6371 | 2 | 5.444 | 4.167 |  |  | Coefficients | Standard Error | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0% | [5 empty] | Coefficients | P-value
25 | Camren | 2025-05-14 14:18 | 2 | 2 | 5 | 2.75 | 94.5312 | 47.2656 | 2 | 6.875 | 2.738 |  | Intercept | 9.5951 | 0.0040505 | 2368.8678 | 1.9494e-139 | 9.587 | 9.6032 | 9.587 | 9.6032 |  |  |  |  | Intercept | 9.5951 | 1.9494e-139
26 | Camren | 2025-05-14 14:28 | 2 | 3 | 1 | 8.4 | 2.9185 | 1.4593 | 2 | 1.208 | 8.392 |  | Op * Laser | 4.9861e-05 | 3.9408e-05 | 1.2652 | 0.21112 | -2.9116e-05 | 0.00012884 | -2.9116e-05 | 0.00012884 |  |  |  |  | Op * Laser | 4.9861e-05 | 0.21112
27 | Camren | 2025-05-14 14:28 | 2 | 3 | 2 | 7.033 | 13.2098 | 6.6049 | 2 | 2.57 | 7.034 |  | Laser^2 | -0.00052632 | 0.00023838 | -2.2079 | 0.031439 | -0.001004 | -4.8598e-05 | -0.001004 | -4.8598e-05 |  |  |  |  | Laser^2 | -0.00052632 | 0.031439
28 | Camren | 2025-05-14 14:28 | 2 | 3 | 3 | 5.628 | 31.7764 | 15.8882 | 2 | 3.986 | 5.614 |  | Op2 | -0.0013599 | 0.0010269 | -1.3243 | 0.19088 | -0.0034179 | 0.00069802 | -0.0034179 | 0.00069802 |  |  |  |  | OP | -0.0013599 | 0.19088
29 | Camren | 2025-05-14 14:28 | 2 | 3 | 4 | 4.256 | 57.4806 | 28.7403 | 2 | 5.361 | 4.251 |  | Laser | -0.99324 | 0.001792 | -554.2518 | 9.6331e-105 | -0.99683 | -0.98965 | -0.99683 | -0.98965 |  |  |  |  | Laser | -0.99324 | 9.6331e-105
30 | Camren | 2025-05-14 14:28 | 2 | 3 | 5 | 2.75 | 94.5588 | 47.2794 | 2 | 6.876 | 2.744
31 | Gabby | 2025-05-19 09:47 | 3 | 1 | 1 | 8.4 | 4.385 | 1.4617 | 3 | 1.209 | 8.388
32 | Gabby | 2025-05-19 09:47 | 3 | 1 | 2 | 7.052 | 19.5075 | 6.5025 | 3 | 2.55 | 7.056
33 | Gabby | 2025-05-19 09:47 | 3 | 1 | 3 | 5.622 | 47.8321 | 15.944 | 3 | 3.993 | 5.618 |  | RESIDUAL OUTPUT
34 | Gabby | 2025-05-19 09:47 | 3 | 1 | 4 | 4.253 | 86.4463 | 28.8154 | 3 | 5.368 | 4.242
35 | Gabby | 2025-05-19 09:47 | 3 | 1 | 5 | 2.75 | 141.7969 | 47.2656 | 3 | 6.875 | 2.74 |  | Observation | Predicted Measure | Residuals
36 | Gabby | 2025-05-19 09:58 | 3 | 2 | 1 | 8.4 | 4.3923 | 1.4641 | 3 | 1.21 | 8.39 |  | 1 | 8.3932 | -0.00019752
37 | Gabby | 2025-05-19 09:58 | 3 | 2 | 2 | 7.066 | 19.3548 | 6.4516 | 3 | 2.54 | 7.049 |  | 2 | 7.0001 | -0.0031142
38 | Gabby | 2025-05-19 09:58 | 3 | 2 | 3 | 5.618 | 47.976 | 15.992 | 3 | 3.999 | 5.62 |  | 3 | 5.5952 | 0.0068076
39 | Gabby | 2025-05-19 09:58 | 3 | 2 | 4 | 4.245 | 86.7041 | 28.9014 | 3 | 5.376 | 4.26 |  | 4 | 4.2583 | -0.0022628
40 | Gabby | 2025-05-19 09:58 | 3 | 2 | 5 | 2.75 | 141.9619 | 47.3206 | 3 | 6.879 | 2.744 |  | 5 | 2.7397 | 0.0023256
41 | Gabby | 2025-05-19 10:08 | 3 | 3 | 1 | 8.4 | 4.385 | 1.4617 | 3 | 1.209 | 8.393 |  | 6 | 8.3942 | 0.00080808
42 | Gabby | 2025-05-19 10:08 | 3 | 3 | 2 | 7.038 | 19.8301 | 6.61 | 3 | 2.571 | 7.042 |  | 7 | 6.9882 | -0.0011654
43 | Gabby | 2025-05-19 10:08 | 3 | 3 | 3 | 5.66 | 46.9498 | 15.6499 | 3 | 3.956 | 5.654 |  | 8 | 5.6141 | -0.0031366
44 | Gabby | 2025-05-19 10:08 | 3 | 3 | 4 | 4.167 | 89.3038 | 29.7679 | 3 | 5.456 | 4.16 |  | 9 | 4.2034 | 0.0046477
45 | Gabby | 2025-05-19 10:08 | 3 | 3 | 5 | 2.75 | 141.9207 | 47.3069 | 3 | 6.878 | 2.743 |  | 10 | 2.7427 | 0.0003262
46 | Dragan | 2025-05-19 10:16 | 4 | 1 | 1 | 8.4 | 5.8371 | 1.4593 | 4 | 1.208 | 8.39 |  | 11 | 8.3942 | 0.0048081
47 | Dragan | 2025-05-19 10:16 | 4 | 1 | 2 | 6.998 | 27.3111 | 6.8278 | 4 | 2.613 | 6.99 |  | 12 | 7.0778 | 0.002222
48 | Dragan | 2025-05-19 10:16 | 4 | 1 | 3 | 5.618 | 63.7762 | 15.944 | 4 | 3.993 | 5.617 |  | 13 | 5.663 | 0.010009
49 | Dragan | 2025-05-19 10:16 | 4 | 1 | 4 | 4.167 | 118.9845 | 29.7461 | 4 | 5.454 | 4.165 |  | 14 | 4.1584 | 0.0025766
50 | Dragan | 2025-05-19 10:16 | 4 | 1 | 5 | 2.75 | 189.5028 | 47.3757 | 4 | 6.883 | 2.745 |  | 15 | 2.7437 | -0.0006736
51 | Dragan | 2025-05-19 10:26 | 4 | 2 | 1 | 8.4 | 5.8467 | 1.4617 | 4 | 1.209 | 8.39 |  | 16 | 8.3899 | 7.8183e-05
52 | Dragan | 2025-05-19 10:26 | 4 | 2 | 2 | 7.067 | 25.8674 | 6.4668 | 4 | 2.543 | 7.06 |  | 17 | 7.0568 | 0.002171
53 | Dragan | 2025-05-19 10:26 | 4 | 2 | 3 | 5.596 | 64.4809 | 16.1202 | 4 | 4.015 | 5.596 |  | 18 | 5.6405 | -0.00048339
54 | Dragan | 2025-05-19 10:26 | 4 | 2 | 4 | 4.227 | 116.4241 | 29.106 | 4 | 5.395 | 4.228 |  | 19 | 4.2464 | -0.0063605
55 | Dragan | 2025-05-19 10:26 | 4 | 2 | 5 | 2.75 | 188.8426 | 47.2106 | 4 | 6.871 | 2.746 |  | 20 | 2.7407 | -0.00067325
56 | Dragan | 2025-05-19 10:51 | 4 | 3 | 1 | 8.4 | 5.8371 | 1.4593 | 4 | 1.208 | 8.392 |  | 21 | 8.3889 | -0.0059275
57 | Dragan | 2025-05-19 10:51 | 4 | 3 | 2 | 7.009 | 27.0192 | 6.7548 | 4 | 2.599 | 7.012 |  | 22 | 7.0598 | -0.0078153
58 | Dragan | 2025-05-19 10:51 | 4 | 3 | 3 | 5.637 | 63.1707 | 15.7927 | 4 | 3.974 | 5.636 |  | 23 | 5.6285 | 0.0074762
59 | Dragan | 2025-05-19 10:51 | 4 | 3 | 4 | 4.214 | 116.9426 | 29.2356 | 4 | 5.407 | 4.205 |  | 24 | 4.1725 | -0.0055193 |  | AGGREGATE DATA SUMMARY OUTPUT
60 | Dragan | 2025-05-19 10:51 | 4 | 3 | 5 | 2.75 | 189.3376 | 47.3344 | 4 | 6.88 | 2.736 |  | 25 | 2.7437 | -0.0056706
Regression Statistics
[17 empty] | Multiple R | 1
[17 empty] | R Square | 0.99999
[17 empty] | Adjusted R Square | 0.99999
[17 empty] | Standard Error | 0.0057959
[17 empty] | Observations | 60
ANOVA
[18 empty] | df | SS | MS | F | Significance F
[17 empty] | Regression | 2 | 239.377 | 119.6885 | 3562984.9083 | 5.4486e-146
[17 empty] | Residual | 57 | 0.0019148 | 3.3592e-05
[17 empty] | Total | 59 | 239.3789
[18 empty] | Coefficients | Standard Error | t Stat | P-value | Lower 95% | Upper 95%
[17 empty] | Intercept | 9.5975 | 0.0023656 | 4057.1181 | 2.4893e-157 | 9.5928 | 9.6023
[17 empty] | OP | -0.00035741 | 0.00066925 | -0.53404 | 0.59539 | -0.0016976 | 0.00098274
[17 empty] | Laser | -0.99649 | 0.00037329 | -2669.4503 | 5.7331e-147 | -0.99724 | -0.99574
RESIDUAL OUTPUT
[17 empty] | Observation | Predicted Measure | Residuals
[17 empty] | 1 | 8.3934 | -0.00041627
[17 empty] | 2 | 6.9983 | -0.001332
[17 empty] | 3 | 5.5933 | 0.0087171
[17 empty] | 4 | 4.258 | -0.0019879
[17 empty] | 5 | 2.7433 | -0.001325
[17 empty] | 6 | 8.3944 | 0.00058724
[17 empty] | 7 | 6.9864 | 0.00062586
[17 empty] | 8 | 5.6122 | -0.0012161
[17 empty] | 9 | 4.2032 | 0.004819
[17 empty] | 10 | 2.7463 | -0.0033145
[17 empty] | 11 | 8.3944 | 0.0045872
[17 empty] | 12 | 7.0761 | 0.0039419
[7 empty] | 0 | 0 | [8 empty] | 13 | 5.661 | 0.011956
[7 empty] | 0 | 0 | [8 empty] | 14 | 4.1583 | 0.002661
[7 empty] | 0 | 0 | [8 empty] | 15 | 2.7473 | -0.004311
[7 empty] | 0 | 0 | [8 empty] | 16 | 8.3911 | -0.0010659
[7 empty] | 0 | 0 | [8 empty] | 17 | 7.0558 | 0.0032291
[7 empty] | 0 | 0 | [8 empty] | 18 | 5.6388 | 0.0012361
[7 empty] | 0 | 0 | [8 empty] | 19 | 4.2457 | -0.0056726
[7 empty] | 0 | 0 | [8 empty] | 20 | 2.743 | -0.0029676
[7 empty] | 0 | 0 | [8 empty] | 21 | 8.3901 | -0.0070694
[7 empty] | 0 | 0 | [8 empty] | 22 | 7.0588 | -0.0067604
[7 empty] | 0 | 0 | [8 empty] | 23 | 5.6268 | 0.0091939
[7 empty] | 0 | 0 | [8 empty] | 24 | 4.1719 | -0.0049325
[7 empty] | 0 | 0 | [8 empty] | 25 | 2.746 | -0.0079571
[7 empty] | 0 | 0 | [8 empty] | 26 | 8.3931 | -0.0010589
[7 empty] | 0 | 0 | [8 empty] | 27 | 7.0358 | -0.0018412
[7 empty] | 0 | 0 | [8 empty] | 28 | 5.6248 | -0.010813
[7 empty] | 0 | 0 | [8 empty] | 29 | 4.2546 | -0.003641
[7 empty] | 0 | 0 | [8 empty] | 30 | 2.745 | -0.00096057
[7 empty] | 0 | 0 | [8 empty] | 31 | 8.3917 | -0.003705
[7 empty] | 0 | 0 | [8 empty] | 32 | 7.0554 | 0.00058646
[7 empty] | 0 | 0 | [8 empty] | 33 | 5.6175 | 0.00051974
[7 empty] | 0 | 0 | [8 empty] | 34 | 4.2473 | -0.0053082
[7 empty] | 0 | 0 | [8 empty] | 35 | 2.7456 | -0.0055997
[7 empty] | 0 | 0 | [8 empty] | 36 | 8.3907 | -0.00070848
[7 empty] | 0 | 0 | [8 empty] | 37 | 7.0654 | -0.016378
[7 empty] | 0 | 0 | [8 empty] | 38 | 5.6115 | 0.0084987
[7 empty] | 0 | 0 | [8 empty] | 39 | 4.2393 | 0.020664
[7 empty] | 0 | 0 | [8 empty] | 40 | 2.7416 | 0.0023863
[7 empty] | 0 | 0 | [8 empty] | 41 | 8.3917 | 0.001295
[7 empty] | 0 | 0 | [8 empty] | 42 | 7.0345 | 0.0075127
[7 empty] | 0 | 0 | [8 empty] | 43 | 5.6544 | -0.00035035
[7 empty] | 0 | 0 | [8 empty] | 44 | 4.1596 | 0.00038279
[7 empty] | 0 | 0 | [8 empty] | 45 | 2.7426 | 0.00038981
[7 empty] | 0 | 0 | [8 empty] | 46 | 8.3923 | -0.0023441
[7 empty] | 0 | 0 | [8 empty] | 47 | 6.9923 | -0.0022773
[7 empty] | 0 | 0 | [8 empty] | 48 | 5.6171 | -0.00012286
[7 empty] | 0 | 0 | [8 empty] | 49 | 4.1613 | 0.0037472
[7 empty] | 0 | 0 | [8 empty] | 50 | 2.7373 | 0.0077297
[7 empty] | 0 | 0 | [8 empty] | 51 | 8.3913 | -0.0013476
[7 empty] | 0 | 0 | [8 empty] | 52 | 7.062 | -0.0020316
[7 empty] | 0 | 0 | [8 empty] | 53 | 5.5952 | 0.0007999
[7 empty] | 0 | 0 | [8 empty] | 54 | 4.22 | 0.0079544
[7 empty] | 0 | 0 | [8 empty] | 55 | 2.7492 | -0.0032282
[7 empty] | 0 | 0 | [8 empty] | 56 | 8.3923 | -0.00034405
[7 empty] | 0 | 0 | [8 empty] | 57 | 7.0062 | 0.0057718
[7 empty] | 0 | 0 | [8 empty] | 58 | 5.6361 | -5.6142e-05
[7 empty] | 0 | 0 | [8 empty] | 59 | 4.2081 | -0.0030877
[7 empty] | 0 | 0 | [8 empty] | 60 | 2.7403 | -0.0042598
* MUI (measurement uncertainty interval) is the expression of the statistical dispersion of the levels measured attributed to the calibration preformed.
** CL = Avg +/- Z * StdDev/N1/2
#39
#37

## Resistance ChA R&R
CAL ID | Operator | Date/Time | OP | Trial | Channel | Slope | Intercept | Slope Dev | Int Dev |  |  |  |  | CL | Z-score
[11 empty] | N | 3 |  | 90% | 1.645 | 2.58
1 | Rick A | 2025-05-07 09:25 | 1 | 2 | Ch A | 112.484 | 0.061 | 0.015667 | 0.020333 |  | NTesters | 3 |  | 96% | 1.96
2 | Rick A | 2025-05-07 13:46 | 1 | 3 | Ch A | 112.459 | 0.109 | 0.0093333 | 0.027667 |  | NReplicates | 4 |  | 99% | 2.58
3 | Rick A | 2025-05-08 10:07 | 2 | 4 | Ch A | 112.462 | 0.074 | 0.0063333 | 0.0073333
[10 empty] | exclude-> |  | Slope | Intercept |  | W | Voltage | LCLmin | LCLmax | MUI* | +/- W
[11 empty] | Average | 112.4683 | 0.081333 |  | 100.08 | 0.889 | 99.9663 | 100.0024 | -0.036152 | 0.018076 | 0.11373
[10 empty] | CL | Std Dev | 0.01365 | 0.024826 |  | 1000.1 | 8.893 | 1000.0001 | 1000.3617 | -0.36165 | 0.18082 | 0.26171
[10 empty] | 99% | UCL | 112.448 | 0.044353
[11 empty] | LCl | 112.4887 | 0.11831 | [5 empty] | ±Tolerance = | 0.1136%
|  |  | ChA | 112.456 | 0.062918 |  | ChB | 113.8686 | -0.14715
|  |  | Ch A | 112.484 | 0.061 |  | Ch B | 113.938 | -1.232
|  |  | Ch A | 112.459 | 0.109 |  | Ch B | 113.933 | -1.202
|  |  | Ch A | 112.462 | 0.074 |  | Ch B | 113.979 | -1.275
CAL ID | Operator | Date/Time | OP | Trial | Channel | Slope | Intercept | Slope Dev | Int Dev
1 | Rick A | 2025-05-07 09:25 | 1 | 2 | Ch A | 112.484 | 0.061 | 0.015667 | 0.020333
2 | Rick A | 2025-05-07 13:46 | 1 | 3 | Ch A | 112.459 | 0.109 | 0.0093333 | 0.027667
3 | Rick A | 2025-05-08 10:07 | 2 | 4 | Ch A | 112.462 | 0.074 | 0.0063333 | 0.0073333
4 | Rick A | 2025-05-07 09:25 | 1 | 2 | Ch B | 113.938 | -1.232 | 0.012 | 0.0043333
5 | Rick A | 2025-05-07 13:46 | 1 | 3 | Ch B | 113.933 | -1.202 | 0.017 | 0.034333
6 | Rick A | 2025-05-08 10:07 | 2 | 4 | Ch B | 113.979 | -1.275 | 0.029 | 0.038667
* MUI (measurement uncertainty interval) is the expression of the statistical dispersion of the levels measured attributed to the calibration preformed.
** CL = Avg +/- Z * StdDev/N1/2

## Resistance ChB R&R
CAL ID | Operator | Date/Time | OP | Trial | Channel | Slope | Intercept | Slope Dev | Int Dev |  |  |  |  | CL | Z-score
[11 empty] | N | 3 |  | 90% | 1.645 | 2.58
2 | Rick A | 2025-05-07 09:25 | 1 | 2 | Ch B | 113.938 | -1.232 | 0.012 | 0.0043333 |  | NTesters | 3 |  | 96% | 1.96
3 | Rick A | 2025-05-07 13:46 | 1 | 3 | Ch B | 113.933 | -1.202 | 0.017 | 0.034333 |  | NReplicates | 4 |  | 99% | 2.58
4 | Rick A | 2025-05-08 10:07 | 2 | 4 | Ch B | 113.979 | -1.275 | 0.029 | 0.038667
[10 empty] | exclude-> |  | Slope | Intercept |  | W | Voltage | LCLmin | LCLmax | MUI* | +/- W
[11 empty] | Average | 113.95 | -1.2363 |  | 100.08 | 0.88 | 100.2429 | 100.3091 | -0.066167 | 0.033084 | 0.22908
[10 empty] | CL | Std Dev | 0.025239 | 0.036692 |  | 1000.1 | 8.789 | 1001.1761 | 1001.837 | -0.66084 | 0.33042 | 1.737
[10 empty] | 99% | UCL | 113.9124 | -1.291
[11 empty] | LCl | 113.9876 | -1.1817 | [5 empty] | ±Tolerance = | 0.2289%
|  |  | ChA | 112.456 | 0.062918 |  | ChB | 113.8686 | -0.14715
|  |  | Ch A | 112.484 | 0.061 |  | Ch B | 113.938 | -1.232
|  |  | Ch A | 112.459 | 0.109 |  | Ch B | 113.933 | -1.202
|  |  | Ch A | 112.462 | 0.074 |  | Ch B | 113.979 | -1.275
* MUI (measurement uncertainty interval) is the expression of the statistical dispersion of the levels measured attributed to the calibration preformed.
** CL = Avg +/- Z * StdDev/N1/2

## Current Source Stability
Fluke 0mA (Eng) | ChA |  | 3-Sigma |  | Reference | MLOI:\$Calibration\Current stability 2025032\Current Stability and Summary 20250401.xlsx
Rref Ω | 100 | 0.01% 15ppm
Samples | 2132
Duration hrs | 117
Average mA | 0.50761
StDev | 6.5041e-06 |  | 1.9512e-05
99.7% +Tolerance | 0.52712
99.7% -Tolerance | 0.4881
Fluke 14mA (Prod) | ChA
Rref Ω | 100 | 0.01% 15ppm
Samples | 28900
Duration hrs | 115
Average mA | 14.7185
StDev | 8.3439e-07 |  | 2.5032e-06
99.7% +Tolerance | 14.721
99.7% -Tolerance | 14.716
DataQ (SF = 0.6, Res 1%) | ChA | ChB
Rref Ω | 100 | 0.01% 15ppm
Samples | 82756
Duration hrs | 115
Average mA | 14.719 | 15.0306
StDev | 1.6614e-05 | 2.0478e-05 | 4.9842e-05 | 6.1434e-05
99.7% +Tolerance | 14.724 | 15.0368
99.7% -Tolerance | 14.714 | 15.0245
PicoScope | ChA | ChB
Rref Ω | 100 | 0.01% 15ppm
Samples | CRASHED!!
Duration hrs
Average mA
StDev
99.7% +Tolerance | 0 | 0
99.7% -Tolerance | 0 | 0
Current Stability Test
This Test of the DataQ with the breadboard PCB version Q01 which has 1% resistors. The current is measured as the voltage accross a 100Ω 0.01% 15ppm TCE reference metal film reference resistors with DataQ and PicoScope. The actual current is measured with the calibrated production Fluke (Fluke 14ma Prod) used as an ammeter on the mA scale.
Engineering Fluke 0mA measurement was with the probes shorted on the mA scale.
DataQ voltage data was colled using WinDaq recording software with a 0.6 scale factor to account for the 1K 1% over 1.5K 1% scaler.
The test was run in the Production Lab from 3/27 to 4/1 for a total of 115hours. The 99.7% tolerance is he +/- 3-sigma range where 99.7% of all measurement will fall within this range as shown below ref: Five Flute - Engineering design review platform for modern hardware teams
The distribution of measurements for the DataQ indicate a resolution of about 50μV

## Switch Point Repeatability
Reference | MLOI:\$ReportFiles\Analysis Summary 250527.xlsx
This is MLOI Q2_01 Run on four different elevate OLS 6 times in succession. The Switch point is determined from the raw data as the time at witch the resistance first enters the switch band.
the Switch band is the area between the upper tolerance of the current level resistor to the lower tolerance of the next resistor value on Draining, the opposite tolerance d for Filling.